Chapter 13: Mechanical Ventilators: General-Use Devices

MULTIPLE CHOICE

1.	Which of the following is (are) used in the measurement of tidal volume on a CareFusion AVEA ventilator?

1. Variable-orifice pneumotachometer

2. Pressure differential pneumotachometer

3. Ultrasonic flow sensor

4. Hot wire flow sensing device

| a. | 1 only | c. | 2 and 3 |
| --- | --- | --- | --- |
| b. | 1 and 4 | d. | 1, 3, and 4 |

ANS:	B

Proximal airway flow sensing with either a variable-orifice pneumotachometer or hot wire flow sensing technology is available for measuring tidal volume at the patient Y connector.

PTS:	1	REF:	Page 429

2.	The Carefusion Avea ventilator is being used during transport; the electrical power source in the ambulance fails. How long will the internal power source allow the ventilator to function?

| a. | About 10 minutes | c. | About 30 minutes |
| --- | --- | --- | --- |
| b. | About 20 minutes | d. | About 40 minutes |

ANS:	C

The internal battery provides about 30 minutes of power when both the ventilator and the compressor are running.

PTS:	1	REF:	Page 420

3.	What mixture of helium:oxygen can be administered with the CareFusion AVEA?

| a. | 60% helium:40% oxygen | c. | 80% helium:20% oxygen |
| --- | --- | --- | --- |
| b. | 70% helium:30% oxygen | d. | 90% helium:10% oxygen |

ANS:	C

The ventilator may be powered with an 80/20 mixture of helium and oxygen.

PTS:	1	REF:	Page 422, Table 13-1

4.	Which of the following pressures can be monitored with the CareFusion AVEA?

1. Esophageal pressure

2. Tracheal pressure

3. Proximal airway pressure

| a. | 3 only | c. | 2 and 3 |
| --- | --- | --- | --- |
| b. | 1 and 2 | d. | 1, 2, and 3 |

ANS:	D

Patient monitoring from the Avea includes esophageal, tracheal, and proximal airway pressure.

PTS:	1	REF:	Page 427

5.	What modes of ventilation on the CareFusion AVEA provide apnea backup ventilation?

1. Volume assist/control

2. Volume synchronized intermittent mandatory ventilation

3. Airway pressure release ventilation/biphasic

4. Continuous positive airway pressure/pressure support ventilation

| a. | 1 only | c. | 2 and 3 |
| --- | --- | --- | --- |
| b. | 1 and 2 | d. | 2, 3, and 4 |

ANS:	D

Apnea backup ventilation is available in all modes in which spontaneous ventilation is available (synchronized intermittent mandatory ventilation, airway pressure release ventilation/biphasic, and continuous positive airway pressure/pressure support ventilation modes).

PTS:	1	REF:	Page 427

6.	How long does the internal battery of the CareFusion AVEA power the ventilator and compressor?

| a. | About 20 minutes | c. | About 40 minutes |
| --- | --- | --- | --- |
| b. | About 30 minutes | d. | About 50 minutes |

ANS:	B

The internal battery provides about 30 minutes of power when both the ventilator and the compressor are running.

PTS:	1	REF:	Page 420

7.	Which of the following respiratory mechanics can be measured with the CareFusion AVEA?

1. Esophageal pressure measurement

2. Peak inspiratory pressure/P100

3. Inflection point

4. Auto-PEEP

| a. | 1 only | c. | 2, 3, and 4 |
| --- | --- | --- | --- |
| b. | 1 and 3 | d. | 1, 2, 3, and 4 |

ANS:	D

Respiratory mechanics that can be measured with the CareFusion AVEA include esophageal pressure measurement, peak inspiratory pressure/P100, inflection point, and auto-PEEP.

PTS:	1	REF:	Page 423

8.	Which of the following on a Dräger Evita XL is considered to be a dual control mode?

| a. | Continuous mandatory ventilation |
| --- | --- |
| b. | Autoflow |
| c. | Synchronized intermittent mandatory ventilation |
| d. | Mandatory minute ventilation |

ANS:	B

Autoflow is a dual control mode of ventilation.

PTS:	1	REF:	Page 429

9.	Which of the following statements describes how a sigh breath is delivered on an Dräger Evita XL ventilator?

| a. | Every 100 breaths the delivered tidal volume is doubled. |
| --- | --- |
| b. | The positive end-expiratory pressure level is increased to the sigh pressure for 2 consecutive breaths. |
| c. | The tidal volume is increased to set sigh volume. |
| d. | The sigh results in doubling the inspiratory time every 100 breaths. |

ANS:	B

Sigh breaths are accomplished on the Dräger EvitaXL ventilator by intermittently increasing the positive end-expiratory pressure level to the set sigh pressure for 2 consecutive breaths in the continuous mandatory ventilation mode.

PTS:	1	REF:	Page 430

10.	What is the available range for tidal volume when the CareFusion AVEA is used in the adult mode?

| a. | 0 to 1500 mL | c. | 100 to 2500 mL |
| --- | --- | --- | --- |
| b. | 100 to 2000 mL | d. | 100 to 3000 mL |

ANS:	C

The range for tidal volume in the adult mode is 100 to 2500 mL.

PTS:	1	REF:	Page 424, Table 13-3

11.	The inspiratory rise time on the CareFusion AVEA is set on one. What is the significance of this setting?

| a. | This results in an inspiratory pause |
| --- | --- |
| b. | There is a slow rise in inspiratory pressure |
| c. | Inspiration is delayed by one tenth of a second |
| d. | There is a rapid rise in inspiratory pressure |

ANS:	D

A setting of one is the most rapid rise in the pressure and a setting of nine is the slowest.

PTS:	1	REF:	Page 424, Table 13-3

12.	Which of the following waveforms is (are) available on the CareFusion AVEA?

1. Sine wave

2. Square wave

3. Decelerating wave

| a. | 3 only | c. | 2 and 3 |
| --- | --- | --- | --- |
| b. | 1 and 3 | d. | 1, 2, and 3 |

ANS:	C

This allows the operator to select either a square wave (constant flow) or a decelerating (descending) ramp flow pattern during volume ventilation. The default is a descending ramp.

PTS:	1	REF:	Page 424, Table 13-3

13.	What is the maximum pressure limit in the CareFusion AVEA?

| a. | 80 cm H2O | c. | 100 cm H2O |
| --- | --- | --- | --- |
| b. | 90 cm H2O | d. | 120 cm H2O |

ANS:	B

Airway pressure release ventilation sets the maximum pressure target 0 to 90 cm H2O (limited to a total peak inspiratory pressure of 90 cm H2O).

PTS:	1	REF:	Page 424, Table 13-3

14.	Which of the following modes of ventilation are available on the CareFusion AVEA?

1. Volume assist/control

2. Pressure assist/control

3. Pressure-regulated volume control (assist/control and synchronized intermittent mandatory ventilation)

4. Volume synchronized intermittent mandatory ventilation

5. Pressure synchronized intermittent mandatory ventilation

6. Airway pressure release ventilation

| a. | 1 and 3 | c. | 1, 2, 4, 5, and 6 |
| --- | --- | --- | --- |
| b. | 2, 3, and 5 | d. | 1, 2, 3, 4, 5, and 6 |

ANS:	D

All of these modes of ventilation are available on the CareFusion AVEA.

PTS:	1	REF:	Page 426, Table 13-4

15.	Which of the following pressures are displayed on the CareFusion AVE?

1. Peak inspiratory pressure (cm H2O)

2. Mean airway pressure (cm H2O)

3. Plateau pressure (cm H2O) if available

4. Positive end-expiratory pressure

| a. | 1 and 2 | c. | 1, 3, and 4 |
| --- | --- | --- | --- |
| b. | 2 and 3 | d. | 1, 2, 3, and 4 |

ANS:	D

All of the above-listed pressures and auto-PEEP are displayed on the ventilator.

PTS:	1	REF:	Page 426, Table 13-5

16.	What is the range for the delivered tidal volume for an adult on the Dräger EvitaXL?

| a. | 0 to 1.0 L | c. | 0 to 1.5 L |
| --- | --- | --- | --- |
| b. | 0 to 1.2 L | d. | 0 to 2.0 L |

ANS:	B

The adult range for tidal volume is 0 to 1.2 liters.

PTS:	1	REF:	Page 429, Table 13-7

17.	What is the range for trigger sensitivity on the Dräger EvitaXL?

| a. | 0.3 to 15 L/min | c. | 1.5 to 25 L/min |
| --- | --- | --- | --- |
| b. | 1.0 to 18 L/min | d. | 2.0 to 30 L/min |

ANS:	A

The trigger sensitivity for the Dräger EvitaXL is 0.3 to 15 L/min.

PTS:	1	REF:	Page 429, Table 13-7

18.	At what point in the ventilatory cycle is auto-PEEP measured on a Dräger EvitaXL?

| a. | Peak inspiration | c. | Mid-expiration |
| --- | --- | --- | --- |
| b. | Inspiratory pause | d. | Expiratory hold |

ANS:	D

The level of auto-PEEP is displayed during an expiratory hold.

PTS:	1	REF:	Page 430, Table 13-8

19.	What is measured by the P0.01 on the Dräger EvitaXL?

| a. | Respiratory muscle strength |
| --- | --- |
| b. | Neurogenic drive |
| c. | Inspiratory flow rate |
| d. | Intrinsic positive end-expiratory pressure |

ANS:	B

The neurogenic drive is measured with the P0.01.

PTS:	1	REF:	Page 430, Table 13-9

20.	What is the range of respiratory rates that can be provided by the Dräger V500?

| a. | 0 to 100 breaths/min | c. | 0 to 150 breaths/min |
| --- | --- | --- | --- |
| b. | 0 to 125 breaths/min | d. | 0 to 200 breaths/min |

ANS:	C

The Dräger V500 can provide respiratory rates between 0 and 150 breaths/min.

PTS:	1	REF:	Page 433, Table 13-11

21.	What positive end-expiratory pressure levels are available on the Dräger V500?

| a. | 0 to 30 cm H2O | c. | 0 to 50 cm H2O |
| --- | --- | --- | --- |
| b. | 0 to 40 cm H2O | d. | 0 to 60 cm H2O |

ANS:	C

The Dräger V500 can provide positive end-expiratory pressure levels of 0 to 50 cm H2O.

PTS:	1	REF:	Page 433, Table 13-11

22.	What modes of ventilation on the Dräger V500 allow for the use of autoflow?

| a. | It can be used only in spontaneous modes of ventilation such as pressure support and continuous positive airway pressure. |
| --- | --- |
| b. | It can be used in pressure-targeted modes such as pressure assist control and pressure synchronized intermittent mandatory ventilation. |
| c. | It can be used in any volume-targeted mode of ventilation, such as continuous mandatory ventilation, synchronized intermittent mandatory ventilation, or mandatory minute ventilation. |
| d. | It can be used with any mode of ventilation. |

ANS:	C

Autoflow is a dual control mode of ventilation. It can be activated when a volume-targeted mode of ventilation is selected, such as continuous mandatory ventilation, synchronized intermittent mandatory ventilation, or maximum voluntary ventilation.

PTS:	1	REF:	Page 433

23.	What mode of ventilation is used on the Dräger V500 to provide a minimum expired volume?

| a. | Minimum minute volume |
| --- | --- |
| b. | Augmented ventilation |
| c. | Mandatory minute volume ventilation |
| d. | Synchronized intermittent mandatory ventilation |

ANS:	C

With mandatory minute volume ventilation, the ventilator provides mandatory breathing only if the patient’s spontaneous breathing is not adequate and drops below the preselected maximum voluntary ventilation setting.

PTS:	1	REF:	Page 430

24.	What is meant by the term “proportional pressure support” on a Dräger V500?

| a. | The level of pressure support is applied in proportion to the patient’s breathing effort. |
| --- | --- |
| b. | The level of pressure support is inversely proportional to the patient’s inspiratory effort. |
| c. | The level of pressure support is applied in proportion to the inspired tidal volume. |
| d. | The level of pressure support is applied in proportion to the work of breathing. |

ANS:	A

Proportional pressure support is a spontaneous breathing mode in which the pressure support is applied in proportion to the patient’s breathing effort.

PTS:	1	REF:	Page 435

25.	What respiratory rate is available on the GE Carestation?

| a. | 1 to 100 breaths/min | c. | 1 to 150 breaths/min |
| --- | --- | --- | --- |
| b. | 1 to 120 breaths/min | d. | 1 to 200 breaths/min |

ANS:	B

The respiratory rate is adjustable from 1 to 120 breaths/min (depending on the mode and breath type).

PTS:	1	REF:	Page 437, Table 13-15

26.	What tidal volume can be delivered by the GE Carestation?

| a. | 10 to 1200 mL | c. | 20 to 2000 mL |
| --- | --- | --- | --- |
| b. | 10 to 1500 mL | d. | 20 to 2500 mL |

ANS:	C

The GE Carestation can deliver tidal volumes between 20 and 2000 mL.

PTS:	1	REF:	Page 437, Table 13-15

27.	What level of pressure support is available on the GE Carestation?

| a. | 0 to 30 cm H2O | c. | 0 to 50 cm H2O |
| --- | --- | --- | --- |
| b. | 0 to 40 cm H2O | d. | 0 to 60 cm H2O |

ANS:	D

The pressure support level on the GE Carestation is between 0 and 60 cm H2O.

PTS:	1	REF:	Page 437, Table 13-15

28.	Which of the following are displayed on the user interface of a GE Carestation?

1. Expired tidal volume

2. Peak inspiratory pressure

3. Rapid shallow breathing index

4. Expired minute volume

| a. | 1 and 3 | c. | 1, 2, and 4 |
| --- | --- | --- | --- |
| b. | 2 and 3 | d. | 1, 2, 3, and 4 |

ANS:	D

The expired tidal volume, peak inspiratory pressure, rapid shallow breathing index, and expired minute volume all are included in the displayed parameters of a GE Carestation.

PTS:	1	REF:	Page 438, Table 13-16

29.	Which of the following can be calculated using the INview Vent Calculation feature found on the GE Carestation?

1. Partial pressure of oxygen in the alveoli

2. Alveolar gas and arterial blood oxygen pressure difference

3. Partial pressure of oxygen in the arteries/fractional inspired oxygen

4. Carbon monoxide

5. Dead space-to-tidal volume ratio

| a. | 1 and 3 | c. | 1, 2, 4, and 5 |
| --- | --- | --- | --- |
| b. | 2, 3, and 5 | d. | 1, 2, 3, 4, and 5 |

ANS:	D

The partial pressure of oxygen in the alveoli, alveolar gas and arterial blood oxygen pressure difference, partial pressure of oxygen in the arteries/fractional inspired oxygen, carbon monoxide, and dead space-to-tidal volume ratio all are calculated with the INview Vent Calculations.

PTS:	1	REF:	Page 438, Table 13-17

30.	Which of the following lung mechanics can be measured by the GE Carestation?

1. P0.1

2. Minute volume

3. Negative inspiratory force

4. Vital capacity

| a. | 1 only | c. | 1, 3, and 4 |
| --- | --- | --- | --- |
| b. | 1 and 2 | d. | 1, 2, 3, and 4 |

ANS:	C

The lung mechanics measured by the GE Carestation include the P0.1, negative inspiratory force, and vital capacity.

PTS:	1	REF:	Page 438, Table 13-17

31.	Which of the following are standard alarms on the GE Carestation?

1. High and low minute volumes

2. High and low respiratory rates

3. High and low airway pressures

4. High and low end-tidal carbon dioxide

| a. | 1 and 3 | c. | 1, 2, and 3 |
| --- | --- | --- | --- |
| b. | 2 and 3 | d. | 1, 2, 3, and 4 |

ANS:	D

The standard alarms on the GE Carestation include high and low minute volumes, high and low respiratory rates, high and low airway pressures, and high and low end-tidal carbon dioxide.

PTS:	1	REF:	Page 438, Table 13-18

32.	Which of the following can be measured with the GE Carestation?

1. Oxygen consumption

2. Carbon dioxide production

3. Ejection fraction

4. Respiratory quotient

5. Energy expenditure

| a. | 1 and 3 | c. | 1, 2, 4, and 5 |
| --- | --- | --- | --- |
| b. | 1, 2, and 4 | d. | 1, 2, 3, 4, and 5 |

ANS:	C

Measurements of oxygen consumption and carbon dioxide production are used to calculate respiratory quotient and energy expenditure.

PTS:	1	REF:	Page 439

33.	What is the range for expiratory trigger sensitivity on the Hamilton G5?

| a. | 5% to 70% of peak inspiratory flow | c. | 10% to 80% of peak inspiratory flow |
| --- | --- | --- | --- |
| b. | 10% to 70% of peak inspiratory flow | d. | 10% to 90% of peak inspiratory flow |

ANS:	A

The expiratory trigger sensitivity on the Hamilton G5 is adjustable from 5% to 70% of the peak inspiratory flow rate.

PTS:	1	REF:	Page 442, Table 13-19

34.	Which of the following flow patterns is (are) available on the Hamilton G5?

1. Decelerating

2. Accelerating

3. Square

4. Sine

| a. | 1 only | c. | 1, 3, and 4 |
| --- | --- | --- | --- |
| b. | 1 and 2 | d. | 1, 2, 3, and 4 |

ANS:	C

The Hamilton G5 offers three inspiratory flow patterns: sine, decelerating, and square waveforms.

PTS:	1	REF:	Page 442, Table 13-19

35.	When an inspiratory pause is used on a Hamilton G5, what is the percentage of the cycle time it can be set for?

| a. | 0 to 50% | c. | 0 to 70% |
| --- | --- | --- | --- |
| b. | 0 to 60% | d. | 0 to 80% |

ANS:	C

When an inspiratory pause is set on the Hamilton G5, it can be set at 0 to 70% of the cycle time.

PTS:	1	REF:	Page 442, Table 13-19

36.	What is the range for setting the level of flow trigger on the Hamilton G5?

| a. | 0.1 to 10 L/min | c. | 0.1 to 20 L/min |
| --- | --- | --- | --- |
| b. | 0.1 to 15 L/min | d. | 0.1 to 25 L/min |

ANS:	B

The effort required on a Hamilton G5 to trigger an assisted breath is adjustable from 0.1 to 15 L/min.

PTS:	1	REF:	Page 442, Table 13-19

37.	What level of pressure support is available on the Hamilton G5?

| a. | 0 to 70 cm H2O | c. | 0 to 90 cm H2O |
| --- | --- | --- | --- |
| b. | 0 to 80 cm H2O | d. | 0 to 100 cm H2O |

ANS:	D

The range for the level of pressure support on the Hamilton G-5 is 0 to 100 cm H2O.

PTS:	1	REF:	Page 442, Table 13-19

38.	Which of the following statements best describes the use of adaptive pressure ventilation on the Hamilton G5?

| a. | The level of pressure support is directly proportional to the peak inspiratory pressure. |
| --- | --- |
| b. | The rise time is adjusted automatically to optimize volume delivery. |
| c. | Volume targeting is provided with pressure-controlled ventilation. |
| d. | Pressure-targeted breaths are provided to minimize risk of barotrauma. |

ANS:	C

Adaptive pressure ventilation provides volume targeting during a pressure-controlled breath. The clinician sets a target tidal volume, and the ventilator adapts to changing lung mechanics by altering the delivery pressure.

PTS:	1	REF:	Page 443

39.	Which of following parameters are measured and displayed on the user interface of a Hamilton G5?

1. Auto-PEEP

2. Peak inspiratory pressure

3. Mean airway pressure

4. Expired minute volume

5. Static compliance

| a. | 2 and 3 | c. | 1, 3, 4, and 5 |
| --- | --- | --- | --- |
| b. | 1, 3, and 4 | d. | 1, 2, 3, 4, and 5 |

ANS:	D

All of the parameters listed are measured and monitored with a Hamilton G5.

PTS:	1	REF:	Page 443, Table 13-20

40.	Which of the following alarms is (are) available on the Hamilton G5?

1. High and low expired minute volume

2. High and low end-tidal carbon dioxide

3. High and low fractional inspired oxygen

4. High and low expired tidal volume

| a. | 1 only | c. | 1, 2, and 4 |
| --- | --- | --- | --- |
| b. | 2 and 3 | d. | 1, 2, 3, and 4 |

ANS:	C

The alarms on the Hamilton G5 include high and low minute volume, high and low end-tidal carbon dioxide, and high and low expired tidal volume.

PTS:	1	REF:	Page 443, Table 13-21

41.	Which of the following is (are) measured by the Hamilton G5?

1. Partial pressure of end-tidal carbon dioxide

2. Airway dead space

3. P0.1

4. Work of breathing caused by endotracheal tube and ventilator circuit

| a. | 1 only | c. | 1, 2, and 4 |
| --- | --- | --- | --- |
| b. | 1 and 3 | d. | 1, 2, 3, and 4 |

ANS:	D

All of the above-listed parameters are measured on the Hamilton G5.

PTS:	1	REF:	Page 444, Table 13-22

42.	What is the range for adjustment of the expiratory sensitivity on the Covidien PB 840?

| a. | 1% to 50% | c. | 1% to 70% |
| --- | --- | --- | --- |
| b. | 1% to 60% | d. | 1% to 80% |

ANS:	D

The expiratory sensitivity on the Covidien PB 840 can be adjusted between 1% and 80%.

PTS:	1	REF:	Page 452, Table 13-27

43.	What is the range for adjustment of the inspiratory pressure on the Covidien PB 840?

| a. | 0 to 80 cm H2O | c. | 10 to 100 cm H2O |
| --- | --- | --- | --- |
| b. | 5 to 90 cm H2O | d. | 10 to 125 cm H2O |

ANS:	B

The inspiratory pressure on the Covidien PB840 can be adjusted between 5 and 90 cm H2O.

PTS:	1	REF:	Page 452, Table 13-27

44.	What range of pressure support is available on the Covidien PB 840?

| a. | 0 to 60 cm H2O | c. | 0 to 80 cm H2O |
| --- | --- | --- | --- |
| b. | 0 to 70 cm H2O | d. | 0 to 90 cm H2O |

ANS:	B

The range for adjustment of pressure support on the Covidien PB 840 is 0 to 70 cm H2O.

PTS:	1	REF:	Page 452, Table 13-27

45.	When proportional assist ventilation is used on the Covidien PB 840, what percentage of support is available?

| a. | 0 to 50% | c. | 10% to 70% |
| --- | --- | --- | --- |
| b. | 5% to 60% | d. | 10% to 80% |

ANS:	D

When proportional assist ventilation is used on the Covidien PB 840, the percentage of support can be adjusted between 10% and 80%.

PTS:	1	REF:	Page 452, Table 13-27

46.	What range of pressure can be displayed on the user interface on the Covidien PB 840?

| a. | –20 to 100 cm H2O | c. | –20 to 150 cm H2O |
| --- | --- | --- | --- |
| b. | –20 to 130 cm H2O | d. | –20 to 180 cm H2O |

ANS:	B

The available range for display of the peak inspiratory pressure, mean airway pressure, and plateau pressure is –20 to 130 cm H2O.

PTS:	1	REF:	Page 453, Table 13-28

47.	Which of the following parameters are displayed on the user interface of the Covidien PB 840?

1. Expired minute volume

2. Peak inspiratory pressure

3. Rapid shallow breathing index

4. Static compliance

5. Intrinsic positive end-expiratory pressure

| a. | 1 and 3 | c. | 1, 2, 4, and 5 |
| --- | --- | --- | --- |
| b. | 2, 3, and 5 | d. | 1, 2, 3, 4, and 5 |

ANS:	D

The display on the user interface of the Covidien PB 840 includes the expired minute volume, peak inspiratory pressure, rapid shallow breathing index, static compliance, and intrinsic positive end-expiratory pressure.

PTS:	1	REF:	Page 453, Table 13-28

48.	Which of the following statements best describes the function of proportional assist ventilation on the Covidien PB 840?

| a. | The percentage of the work done by the patient is adjusted by the ventilator monitoring the work of breathing. |
| --- | --- |
| b. | The percentage of the work done by the ventilator is automatically reduced as the compliance and resistance is improving. |
| c. | The operator sets a percentage of the work of breathing that is done by the ventilator; the remainder of the work of breathing is done by the patient. |
| d. | The ventilator adjusts the amount of work done by the ventilator based on continuous monitoring of the partial pressure of end-tidal carbon dioxide and the oxygen saturation measured by pulse oximeter. |

ANS:	C

The proportional assist breath targets a pressure based on a selected percentage of support set by the operator and the flow and volume reading from the patient.

PTS:	1	REF:	Page 452

49.	Which of the following statements best describes the volume support mode on the Covidien    PB 840?

| a. | The ventilator ensures that a minimum of 80% of the set tidal volume is delivered with each breath. |
| --- | --- |
| b. | The ventilator delivers volume-targeted breaths with each inspiratory effort. |
| c. | The ventilator delivers pressure-limited breaths; the pressure is adjusted to ensure that the set tidal volume is delivered. |
| d. | The ventilator delivers pressure-targeted breaths, and the respiratory rate is adjusted to ensure delivery of a preset minute volume. |

ANS:	C

Volume support is pressure support ventilation with a volume target. In volume support, breaths are patient triggered (pressure or flow), pressure limited, and flow cycled. The ventilator automatically adjusts the pressure limit to achieve the set tidal volume.

PTS:	1	REF:	Page 452

50.	Which of the following can be monitored with the Covidien PB 840?

1. Leak flow and volume

2. Dynamic compliance and airway resistance

3. Work of breathing during proportional assist ventilation

4. Peak spontaneous flow

| a. | 1 and 3 | c. | 1, 3, and 4 |
| --- | --- | --- | --- |
| b. | 2 and 3 | d. | 1, 2, 3, and 4 |

ANS:	D

Advanced monitoring capabilities of the Covidien PB 840 include the ability to monitor leak flow and volume, dynamic compliance and airway resistance, work of breathing during proportional assist ventilation, and peak spontaneous flow.

PTS:	1	REF:	Page 453, Table 13-29

51.	Which of the following alarms are used with the Covidien PB 840 ventilator?

1. Apnea interval

2. High and low circuit pressure

3. High and low exhaled minute volume

4. High intrinsic positive end-expiratory pressure

5. High and low spontaneous tidal volume

| a. | 1 and 3 | c. | 1, 2, 3, and 5 |
| --- | --- | --- | --- |
| b. | 2, 3, and 4 | d. | 1, 2, 3, 4, and 5 |

ANS:	C

The Covidien PB 840 includes an apnea interval alarm, high and low circuit pressure alarm, high and low exhaled minute volume, and high and low spontaneous tidal volume alarm.

PTS:	1	REF:	Page 453, Table 13-30

52.	What is the available range for the respiratory rate when the Maquet Servoi is used in the continuous mandatory ventilation mode?

| a. | 1 to 60 breaths/min | c. | 4 to 100 breaths/min |
| --- | --- | --- | --- |
| b. | 1 to 80 breaths/min | d. | 4 to 125 breaths/min |

ANS:	C

The respiratory rate is adjustable from 1 to 60 breaths/min when the synchronized intermittent mandatory ventilation mode is used and from 4 to 100 breaths/min when the continuous mandatory ventilation mode is used.

PTS:	1	REF:	Page 456, Table 13-32

53.	Which of the following is the available range for pressure support on the Maquet Servoi?

| a. | 0 to (120 cm H2O – positive end-expiratory pressure) |
| --- | --- |
| b. | 0 to (100 cm H2O – positive end-expiratory pressure) |
| c. | 0 to (80 cm H2O – positive end-expiratory pressure) |
| d. | 0 to (60 cm H2O – positive end-expiratory pressure) |

ANS:	A

The range for pressure support on the Maquet Servoi is 0 to (120 cm H2O – positive end-expiratory pressure).

PTS:	1	REF:	Page 456, Table 13-32

54.	What is the range for the tidal volume on the Maquet Servoi?

| a. | 0 to 1500 mL | c. | 100 to 2000 mL |
| --- | --- | --- | --- |
| b. | 100 to 1500 mL | d. | 100 to 2500 mL |

ANS:	C

The range for tidal volume for the Maquet Servoi is 100 to 2000 mL.

PTS:	1	REF:	Page 456, Table 13-32

55.	Which of the following parameters is (are) displayed on the user interface of the Maquet Servoi?

1. Expired minute volume

2. Inspiratory tidal volume

3. Intrinsic positive end-expiratory pressure

4. Shallow breathing index

| a. | 2 only | c. | 1, 2, and 4 |
| --- | --- | --- | --- |
| b. | 1 and 3 | d. | 1, 2, 3, and 4 |

ANS:	C

The display of the Maquet Servoi includes the expired minute volume, inspiratory tidal volume, and shallow breathing index. The total positive end-expiratory pressure is displayed; this includes the intrinsic positive end-expiratory pressure and the applied positive end-expiratory pressure.

PTS:	1	REF:	Page 457, Table 13-33

56.	Which of the following statements best describes automode on the Maquet Servoi?

| a. | The ventilator adjusts the delivered tidal volume based on the partial pressure of end-tidal carbon dioxide and the oxygen saturation measured by pulse oximeter. |
| --- | --- |
| b. | The ventilator selects the optimal tidal volume and rate based on the compliance and resistance. |
| c. | The ventilator automatically shifts between controlled ventilation and supported or spontaneous ventilation based on the needs of the patient. |
| d. | The ventilator adjusts the delivered tidal volume based on the measured work of breathing. |

ANS:	C

Automode allows the ventilator to shift automatically between controlled ventilation, supported ventilation, and spontaneous ventilation based on the effort sensed from the patient.

PTS:	1	REF:	Page 457

57.	Which of the following ventilators can be used during magnetic resonance imaging?

| a. | CareFusion AVEA | c. | Maquet Servoi |
| --- | --- | --- | --- |
| b. | Covidien PB 840 | d. | Hamilton G-5 |

ANS:	C

The Maquet Servoi can be adapted for use in a magnetic resonance imaging environment using up to a 3-Tesla magnet; this makes an intensive care ventilator available in the magnetic resonance imaging suite.

PTS:	1	REF:	Page 457

58.	Which of the following is a difference between the Maquet Servoi and the Maquet Servos?

| a. | The ServoS is able to deliver larger tidal volumes. |
| --- | --- |
| b. | The ServoS is able to deliver a mixture of helium and oxygen. |
| c. | The ServoS cannot be used to manage neonatal patients, whereas the Maquet Servoi can be used with neonatal patients. |
| d. | The ServoS can be adapted to deliver oscillatory ventilation. |

ANS:	C

The Maquet Servoi and the Maquet Servos are identical in functional design and user operation. The Maquet Servos can be used for pediatric and adult patients but not for neonatal patients.

PTS:	1	REF:	Page 454

59.	Which of the following statements best describes the application of the Open Lung Tool on the Maquet Servoi?

| a. | It is used to measure the level of the patient’s work of breathing. |
| --- | --- |
| b. | It is used to prevent overdistention and collapse of the alveolus. |
| c. | It is used to measure the static compliance and the airway resistance. |
| d. | It can be used to select the optimal minute ventilation for the patient. |

ANS:	B

The Open Lung Tool allows the operator to select any method of evaluating the pressure of overdistention and the pressure that occurs when the lung unit collapses.

PTS:	1	REF:	Page 457

60.	Which of the following can be monitored with the Maquet Servoi?

1. Inspiratory and expiratory resistance

2. Patient work of breathing

3. Work of breathing associated with the ventilator

4. Electrical activity of the diaphragm

| a. | 1 and 3 | c. | 1, 3, and 4 |
| --- | --- | --- | --- |
| b. | 2 and 3 | d. | 1, 2, 3, and 4 |

ANS:	D

Advanced monitoring capabilities of the Maquet Servoi include measurement of the inspiratory and expiratory resistance, patient work of breathing, the work of breathing associated with the ventilator, and electrical activity of the diaphragm.

PTS:	1	REF:	Page 456, Table 13-32