Chapter 06: Humidity and Aerosol Therapy

MULTIPLE CHOICE

1.	Which of the following statements is true?

| a. | Water that exists in the form of individual molecules in a gaseous state is humidity. |
| --- | --- |
| b. | Water that exists in the form of individual molecules is an aerosol. |
| c. | Water that exists in the form of particulate matter is relative humidity (RH). |
| d. | Water that exerts no pressure is a vapor. |

ANS:	A

Humidity is water that exists in the form of individual molecules in the vaporous or gaseous state. An aerosol is a suspension of solid or liquid particles in a gas. Water vapor does exert pressure, although the pressure may vary.

PTS:	1	REF:	Page 159

2.	Which of the following statements is true?

| a. | Fog and mist are considered water vapor. |
| --- | --- |
| b. | Medical aerosols are much smaller than molecules of water. |
| c. | Medical aerosols range from 0.2 m to 50 m. |
| d. | All of the above are true. |

ANS:	C

Medical aerosol must deposit anywhere from the upper airways to the lower airways; therefore, it must be the appropriate size for the target site, with a range of 0.2 m to 50 m. Fog and mist are considered aerosols, not water vapor. Medical aerosols need to be larger than water vapor molecules—otherwise they would be too stable to deposit at their target site.

PTS:	1	REF:	Page 159

3.	A patient is receiving albuterol through a small-volume nebulizer. Where in the body will the greatest effect take place?

| a. | Localized effect in the lungs |
| --- | --- |
| b. | Systemic effect |
| c. | Local and systemic effect |
| d. | Depends on the size, temperature, and volume of the aerosol |

ANS:	D

Aerosols have the potential to cause both local and systemic effects. However, these effects largely depend on the drug concentration delivered, particle size, temperature, and volume of the aerosol.

PTS:	1	REF:	Page 181

4.	Which of the following is true concerning water vapor?

| a. | It is a true gas. |
| --- | --- |
| b. | It has a constant pressure. |
| c. | It exerts a partial pressure. |
| d. | It exists above its critical temperature. |

ANS:	C

Water vapor exerts a pressure (PH2O) that results from the continuous, random movement of water molecules. It is not a true gas and does not have a constant pressure. As the temperature of a gas increases, water vapor pressure increases and water molecules leave the liquid state through evaporation.

PTS:	1	REF:	Page 159

5.	A water molecule at the surface contains enough kinetic activity to enter the vapor state. This is the definition of which of the following?

| a. | Evaporation | c. | The boiling point |
| --- | --- | --- | --- |
| b. | Vaporization | d. | The critical point |

ANS:	A

Evaporation occurs when liquid molecules near the surface contain enough kinetic energy to break free and enter a vapor state without the addition of heat. Vaporization is the energy required to vaporize a liquid. The boiling point is when the water is heated to a point at which molecules leave the liquid state and produce water vapor.

PTS:	1	REF:	Page 159

6.	When absolute humidity (AH) is constant and the temperature of the gas increases, the:

| a. | gas is saturated. | c. | water content increases. |
| --- | --- | --- | --- |
| b. | capacity decreases. | d. | relative humidity decreases. |

ANS:	D

Relative humidity decreases when the absolute humidity is held constant and the temperature of the gas is increased because the warmer gas has a higher capacity. Mathematically, when the denominator of a ratio is increased, the whole number will decrease.

PTS:	1	REF:	Page 159

7.	What is the relative humidity when the absolute humidity is 15 mg H2O/L and the capacity of a gas is 20.6 mg H2O/L?

| a. | 15% | c. | 30% |
| --- | --- | --- | --- |
| b. | 20% | d. | 72% |

ANS:	D

RH = (AH/capacity)  100%. (15 mg/L/20.6 mg/L)  100% = 72%.

PTS:	1	REF:	Page 159, Box 6-1

8.	The relative humidity is 40% and the capacity is 32 mg H2O/L. The absolute humidity of this gas is _____ mm Hg.

| a. | 1.3 | c. | 60 |
| --- | --- | --- | --- |
| b. | 12.8 | d. | 128 |

ANS:	B

RH = (AH/capacity)  100%. [40% = 0.4]. 0.4 = x/32. 32  0.4 = 12.8 mg H2O/L.

PTS:	1	REF:	Page 159, Box 6-1

9.	At 20° C, water vapor pressure is measured as 15 mm Hg. If the maximum partial pressure at this temperature is 17.5 mm Hg, the relative humidity is:

| a. | 30%. | c. | 75%. |
| --- | --- | --- | --- |
| b. | 50%. | d. | 85%. |

ANS:	D

15/17.5  100% = 85%.

PTS:	1	REF:	Page 159, Box 6-1

10.	If its relative humidity is 50% and its absolute humidity is 16 mg H2O/L, the water vapor capacity of a gas is _____ mg H2O/L.

| a. | 0.8 | c. | 8 |
| --- | --- | --- | --- |
| b. | 3.2 | d. | 32 |

ANS:	D

0.5 = 16/X.

0.5X = 16.

X = 16/0.5.

X = 32 mg H2O.

PTS:	1	REF:	Page 159, Box 6-1

11.	At 37° C, water vapor exerts a partial pressure of ____ mm Hg and contains _____ mg H2O/L.

| a. | 43.9; 47 | c. | 44; 44 |
| --- | --- | --- | --- |
| b. | 47; 43.9 | d. | 100; 100 |

ANS:	B

At sea level, a gas at body temperature (37° C), at a pressure of 47 mm Hg, and at a saturation of 100% will contain 43.9 mg of water per liter of gas.

PTS:	1	REF:	Page 159, Box 6-1

12.	Condensation will form in large-bore corrugated tubing during the delivery of bland aerosol because the:

| a. | gas is warmed by ambient air. | c. | gas is cooled by ambient air. |
| --- | --- | --- | --- |
| b. | corrugated tubing is too short. | d. | particles are susceptible to gravity. |

ANS:	D

The liquid particles suspended in gas grow larger and heavier, condense, and then fall out because large particles are more susceptible to gravity than small particles. If the ambient air is warmer than the gas carrying the aerosol, the gas will warm, increasing its capacity to hold water, thereby preventing rain out in the tubing. The shorter the tubing, the less time there is for the rain out to occur before the aerosol reaches the patient. The ambient air in the hospital setting is usually at or warmer than the temperature of the gas carrying the aerosol. So, the gas will not cool off as it flows to the patient unless it is a heated aerosol.

PTS:	1	REF:	Page 160

13.	A pulse oximeter on a patient with a tracheostomy shows 89% saturation. The patient has a heated humidifier with a 40% air-entrainment tracheostomy collar. A rapid assessment shows no secretion production and equal breath sounds. Which of the following is the most appropriate first step?

| a. | Increase the oxygen delivered through the flowmeter. |
| --- | --- |
| b. | Increase the air-entrainment device to 60%. |
| c. | Apply positive pressure with bag-valve ventilation. |
| d. | Check to see whether water is collected in the gas-delivery tube. |

ANS:	D

The condensation collected in the gas-delivery tube must be removed because it becomes an obstruction that disrupts gas and humidity delivery to the patient. The other choices disregard the simplest step, which should be taken first.

PTS:	1	REF:	Page 160

14.	Which of the following is true about the physiologic control of heat and moisture exchange by the body?

| a. | The nose, oropharynx, and hypopharynx add heat and humidity during exhalation. |
| --- | --- |
| b. | The nose, oropharynx, and hypopharynx cool gas and reclaim water during inhalation. |
| c. | The inability of the nasal mucosa to reabsorb excessive amounts of exhaled condensation in cold weather causes a “runny nose.” |
| d. | The mouth is more efficient at reclaiming heat and water during exhalation than is the nose. |

ANS:	C

A runny nose in a cold environment is caused by the inability of the nasal mucosa to reabsorb excessive amounts of exhaled gas condensed in water. The upper airways add heat and humidity to inspired gases during inhalation, and cool and reclaim water during exhalation. During exhalation, the mouth is much less efficient than the nose in reclaiming heat and water.

PTS:	1	REF:	Page 160

15.	The isothermic saturation boundary is located just below the:

| a. | carina. | c. | lobar bronchi. |
| --- | --- | --- | --- |
| b. | nasopharynx. | d. | mainstem bronchi. |

ANS:	A

The isothermic saturation boundary is the point, typically 5 cm below the carina, at which body temperature and pressure, saturated with water vapor (BTPS) is reached. The nasopharynx is above the carina.

PTS:	1	REF:	Page 160, Figure 6-1

16.	Which of the following statements about the isothermic saturation boundary is true?

| a. | Exercise causes the isothermic saturation boundary to shift higher in the trachea. |
| --- | --- |
| b. | At 2 cm above the carina, inhaled gas reaches BTPS. |
| c. | Inhalation through the mouth shifts the isothermic saturation boundary downward. |
| d. | Below the isothermic saturation boundary, temperature and relative humidity are variables. |

ANS:	C

At the isothermic saturation boundary inhaled gas reaches BTPS, which is below the carina. The temperature and relative humidity are constant after that point. Because exercise causes an increase in the volume, rate, and type of breathing (i.e., breathing through the mouth), the isothermic saturation boundary would shift downward.

PTS:	1	REF:	Pages 160-161

17.	Which of the following are true about the function of the upper airways?

1. They cool inspired air.

2. They filter out secretions.

3. They humidify inspired air.

4. They keep the nasal mucosa moist.

| a. | 1, 2, and 3 | c. | 2 and 4 |
| --- | --- | --- | --- |
| b. | 1 and 3 | d. | 3 and 4 |

ANS:	D

The upper airway is responsible for humidifying inspired air. The upper airways add heat and humidity to inspired gases during inhalation. The nasal mucosa is kept moist by secretions from the mucous glands. Humidity and warmth are captured by the nasal mucosa, and exhaled air is cooled.

PTS:	1	REF:	Page 160

18.	A large humidity deficit will cause:

| a. | a nonproductive cough. | c. | increased movement of mucus. |
| --- | --- | --- | --- |
| b. | increased cilia function. | d. | decreased airway resistance. |

ANS:	A

When the airways are exposed to dry, cold air, ciliary motility is reduced; the airways become more irritable; and because mucous production is increased, pulmonary secretions become thick and inspissated, with increased airway resistance and a nonproductive cough.

PTS:	1	REF:	Page 161

19.	The structure that has a large surface area and therefore increases contact between inspired air and mucosa is the:

| a. | nose. | c. | larynx. |
| --- | --- | --- | --- |
| b. | carina. | d. | epiglottis. |

ANS:	A

In the nose, the tortuous path of the turbinates increases contact between the inspired air and the mucosa.

PTS:	1	REF:	Page 160

20.	Inspired air is humidified and warmed through what two physical processes?

1. Condensation

2. Evaporation

3. Convection

4. Conduction

| a. | 1 and 2 | c. | 2 and 3 |
| --- | --- | --- | --- |
| b. | 1 and 4 | d. | 3 and 4 |

ANS:	C

As the inspired air enters the nose, it is warmed by convection, not by conduction, and picks up water vapor from the moist mucosal lining by evaporation, not by condensation.

PTS:	1	REF:	Page 160

21.	A high-flow nasal cannula set at 15 L has been running for a while without water. Which of the following could have happened?

| a. | The cilia function increases. |
| --- | --- |
| b. | Turbulent convection stops. |
| c. | Airway resistance decreases. |
| d. | The isothermic saturation boundary shifts. |

ANS:	D

The isothermic saturation boundary shifts deeper into the lungs when a person inhales through cold, dry air. In this case, the cilia function decreases and secretions are produced, increasing airway resistance. The turbulent convection increases contact between the inspired air and the mucosa.

PTS:	1	REF:	Page 160

22.	The relative humidity of an inspired gas is greater than 60% of BTPS. Which of the following statements is true?

| a. | Production of mucus in the airways will increase. |
| --- | --- |
| b. | The work of breathing will increase. |
| c. | The airways will function normally. |
| d. | Secretions will become thick and dehydrated. |

ANS:	C

When the relative humidity of inspired gas is greater than 60% of BTPS conditions, no injury is believed to occur in normal lungs. The production of mucus will increase, work of breathing will increase, and secretions will become thick and dehydrated when the relative humidity is less than 60% of BTPS.

PTS:	1	REF:	Page 161

23.	Indications for the use of warm, humidified gas are which of the following?

1. Patients with hypothermia

2. Patients with croup

3. Patients with bronchospasm

4. Patients with postextubation edema

| a. | 1, 2, and 3 | c. | 2 and 4 |
| --- | --- | --- | --- |
| b. | 1 and 3 | d. | 1, 2, 3, and 4 |

ANS:	B

Warm, humidified gases are used to raise core temperatures of patients with hypothermia to prevent intraoperative hypothermia, and to alleviate bronchospasm in patients with reactive airways, in addition to maintaining normal humidity levels in the airway. Cool, humidified gas is often used in conjunction with a bland aerosol delivery to treat upper airway inflammation resulting from croup, epiglottitis, and postextubation edema.

PTS:	1	REF:	Page 161

24.	A patient with a nasal cannula set to deliver 8 L/min of oxygen would be expected to report which of the following?

1. Dry, nonproductive cough

2. Increased work of breathing

3. Substernal pain

4. Thick, dehydrated secretions

| a. | 1, 2, and 3 | c. | 2 and 4 |
| --- | --- | --- | --- |
| b. | 1 and 3 | d. | 1, 2, 3, and 4 |

ANS:	D

Dry, nonproductive cough; increased work of breathing; substernal pain; and thick, dehydrated secretions are symptoms of inadequate airway humidification. See Box 4-2.

PTS:	1	REF:	Page 161, Box 6-2

25.	Which of the following statements about humidifiers is appropriate?

| a. | The warmer the gas, the less water vapor it can hold. |
| --- | --- |
| b. | Evaporation occurs more rapidly with a large water-to-gas surface area. |
| c. | High peak flows improve the efficacy of a humidifier. |
| d. | The mass of the water is inversely proportional to its capacity to transfer heat. |

ANS:	B

Larger surface areas will increase the amount of evaporation that will occur. The warmer the gas, the more water it can hold. High flows through a humidifier decrease its efficiency, because there is less time for contact between the gas and water. According to the principles of thermodynamics, the mass of the water is directly proportional to its capacity to transfer heat. See also Box 6-4.

PTS:	1	REF:	Page 161, Box 6-4

26.	Which of the following humidifiers is classified as passive?

| a. | Bubble humidifier | c. | Passover humidifier |
| --- | --- | --- | --- |
| b. | Large-volume nebulizer | d. | Heat and moisture exchanger |

ANS:	D

A heat and moisture exchanger is a passive humidifier. All the others are active humidifiers.

PTS:	1	REF:	Page 176

27.	When the temperature is 37° C and the relative humidity is 35%, the absolute humidity is _____ mg H2O/L.

| a. | 11.9 | c. | 16.5 |
| --- | --- | --- | --- |
| b. | 15.4 | d. | 79.5 |

ANS:	B

RH = [content (AH)/capacity]  100%. (35% = 0.35). 0.35 = x/44 = 15.4 mg H2O/L.

PTS:	1	REF:	Page 159, Box 6-1

28.	If the relative humidity of a gas is 55% and the capacity is 18 mg/L, the absolute humidity is _____ mg H2O/L.

| a. | 10 | c. | 22.5 |
| --- | --- | --- | --- |
| b. | 12 | d. | 25.2 |

ANS:	A

RH = (content [AH)/capacity].

0.55 = x/18 mg H2O/L.

x = 10 mg H2O/L.

PTS:	1	REF:	Page 159, Box 6-1

29.	What is the relative humidity of air at 37° C if the actual water vapor pressure is 33 mm Hg?

| a. | 89% | c. | 70% |
| --- | --- | --- | --- |
| b. | 75% | d. | 25% |

ANS:	C

RH = [content(AH)/capacity]  100%.

(33 mm Hg/47 mm Hg [water vapor pressure at body temperature])  100% = 70%.

PTS:	1	REF:	Page 159, Box 6-1

30.	The actual water vapor pressure at 37° C when the relative humidity is 50% is _____ mm Hg.

| a. | 18.5 | c. | 22 |
| --- | --- | --- | --- |
| b. | 23.5 | d. | 45 |

ANS:	B

RH = [content (AH)/capacity]  100%.

0.5 = x/44 mm Hg = 23.5 mm Hg.

PTS:	1	REF:	Page 159, Box 6-1

31.	If the actual water vapor pressure is 40 mm Hg and the capacity is 55 mm Hg, the relative humidity is:

| a. | 88%. | c. | 80%. |
| --- | --- | --- | --- |
| b. | 85%. | d. | 35%. |

ANS:	C

RH = [content (AH)/capacity]  100%.

(44/55)  100% = 80%.

PTS:	1	REF:	Page 150, Box 6-1

32.	The efficiency of a bubble humidifier is affected by which of the following?

| a. | Gas flow rate | c. | Aerosol production |
| --- | --- | --- | --- |
| b. | Operating pressure | d. | Oxygen concentration |

ANS:	A

As gas flow increases, the reservoir cools by as much as 10° C, thus limiting its effectiveness to provide humidity at flow rates higher than 10 L/min. An increased gas flow rate decreases the amount of contact time between the gas and the water. Increasing the height of the water column above the gas outlet increases the humidity content of the bubbles by allowing for the longer contact time.

PTS:	1	REF:	Page 161

33.	A patient reports a dry nose and throat. He is receiving 6 L of O2 through a nasal cannula. Which of the following is most appropriate to address the patient’s dry nose and throat?

| a. | Decrease the oxygen flow rate to 3 L. |
| --- | --- |
| b. | Add a low-flow bubble humidifier. |
| c. | Suggest an aerosol treatment with a mucolytic agent. |
| d. | Take no action. |

ANS:	B

If a nasal cannula is set to deliver more than 5 L/min, humidification will be necessary. A low-flow bubble humidifier humidifies the gas traveling to the patient through a nasal cannula with approximately 40% to 50% of the relative humidity at standard room temperatures. See Figure 6-4. The other options are not applicable.

PTS:	1	REF:	Page 163, Figure 6-4

34.	A low-flow bubble humidifier is running at 4 L/min and is not whistling. Which action should be taken at this time?

| a. | No action should be taken because this is normal. |
| --- | --- |
| b. | Increase the flow to make it whistle. |
| c. | Replace the spring-loaded valve. |
| d. | Change the unit because there is a leak. |

ANS:	A

The humidifier is functioning normally. The pressure-relief valve releases high pressure (>2 psi) to prevent bursting of the humidifier bottle when there is an obstruction creating a whistling sound. Leaks are not detected by the pressure-relief valve.

PTS:	1	REF:	Page 162

35.	Which of the following is a true statement about the delivery of warm, humidified gas to the lungs?

| a. | The gas delivered to the nose or mouth contains 95% relative humidity and has an absolute humidity of 28 mg/L to 34 mg/L at 29° C to 32° C. |
| --- | --- |
| b. | The gas delivered to the hypopharynx contains 50% relative humidity and has an absolute humidity level of 10 mg/L at 22° C. |
| c. | The gas delivered to the midtrachea contains 50% relative humidity and has an absolute humidity level of 36 mg/L to 40 mg/L at 31° C to 35° C. |
| d. | The gas delivered to the midtrachea contains 100% relative humidity and has an absolute humidity level of 36 mg/L to 40 mg/L at 31° C to 35° C. |

ANS:	D

The gas delivered to the midtrachea contains 100% relative humidity and has an absolute humidity level of 36 mg/L to 40 mg/L at 31° C to 35° C. The gas delivered to the hypopharynx contains 95% relative humidity and has an absolute humidity level of 28 mg/L to 34 mg/L at 29° C to 32° C. The gas delivered to the nose or mouth contains 50% relative humidity and has an absolute humidity level of 10 mg/L at 22° C. See Box 6-3.

PTS:	1	REF:	Page 161, Box 6-3

36.	The efficiency of a bubble humidifier depends on which of the following?

1. Presence of a diffuser

2. Water level in the humidifier

3. Surface area between the gas and water

4. Length of the oxygen tubing being used

| a. | 1 only | c. | 3 and 4 only |
| --- | --- | --- | --- |
| b. | 2 and 4 only | d. | 1, 2, and 3 only |

ANS:	D

The effectiveness of bubble humidifiers depends of the presence of a diffuser, the water level in the humidifier, and the surface area between the gas and water. The length of the oxygen tubing being used will not affect the efficiency of the bubble humidifier.

PTS:	1	REF:	Page 161, Box 6-4

37.	A nonrebreathing mask has been replaced by a high-flow nasal cannula humidifier set at 15 L/min. The benefit of this device to a patient is that it:

| a. | decreases airway resistance. |
| --- | --- |
| b. | delivers medication to the airways. |
| c. | prevents an increased humidity deficit. |
| d. | dilutes thick secretions for easy removal. |

ANS:	C

The high-flow nasal cannula humidifier will prevent an increase in humidity deficit because it uses a special bubble humidifier that can accommodate the rapid flow rates and still humidify the gas.

PTS:	1	REF:	Page 163

38.	The respiratory therapist has been called because the nasal cannula with humidifier delivering 4 L/min is making a high-pitched whistle. The most probable cause of the problem is that the:

| a. | humidifier is leaking. | c. | oxygen tubing is too long. |
| --- | --- | --- | --- |
| b. | oxygen tubing is kinked. | d. | nasal prong is not in place. |

ANS:	B

Humidifier pop-off valves typically provide both an audible alarm and a visible alarm; the humidifier automatically resumes normal operation once the excess pressure is vented.

PTS:	1	REF:	Page 162

39.	Which of the following statements is true concerning bland aerosols?

| a. | Bland aerosols should not be delivered to patients with tracheostomies. |
| --- | --- |
| b. | Ultrasonic nebulizers are inefficient at producing bland aerosols. |
| c. | Sterile water and hypotonic, isotonic, and hypertonic solutions can be used. |
| d. | Humidifiers must be heated to deliver bland aerosols. |

ANS:	C

Sterile water and hypotonic, hypertonic, and isotonic solutions can all be used to create bland aerosols. If the patient has a tracheostomy, this requires the addition of a device that can produce high humidity. In most cases, this device would be the large-volume nebulizer, which produces bland aerosols. Ultrasonic nebulizers are very efficient at producing bland aerosols. Large-volume nebulizers (LVNs) do not need heat to produce bland aerosols.

PTS:	1	REF:	Page 200, Clinical Practice Guideline 6-2

40.	Appropriate use of an ultrasonic nebulizer will aid which of the following?

| a. | Ribavirin delivery | c. | Croup tent use |
| --- | --- | --- | --- |
| b. | Sputum induction | d. | Pentamidine aerosolization |

ANS:	B

Ultrasonic nebulizers generate bland aerosols and aid the delivery of sterile water and hypotonic, isotonic, or hypertonic saline aerosols. They are very effective when used to induce sputum production for laboratory testing. Ribavirin is aerosolized in the small-particle aerosol generator model 2 (SPAG-2). Pentamidine is aerosolized in the Circulaire or Respirgard type of nebulizers that have filters to remove the aerosol before it is emitted into the environment.

PTS:	1	REF:	Page 206

41.	A respiratory therapist needs to deliver hypotonic solution to a patient through a large-volume jet nebulizer. The gas flow setting should be ___ L/min to ___ L/min.

| a. | 2; 4 | c. | 6; 8 |
| --- | --- | --- | --- |
| b. | 4; 6 | d. | 6; 15 |

ANS:	D

Jet nebulizers typically operate at a flow of 6 L/min to 15 L/min.

PTS:	1	REF:	Page 164

42.	Which of the following statements regarding large-volume jet nebulizers is true?

| a. | Gas is brought down the siphon tube into the liquid. |
| --- | --- |
| b. | Lateral pressure increases around the air-entrainment ports. |
| c. | Baffling removes the larger particles from the generated aerosol. |
| d. | Inspiratory flow can be increased by closing the air-entrainment ports. |

ANS:	C

The resulting low pressure generated at the jet draws fluid from the reservoir up to the top of a siphon tube, where it is sheared off and is carried in the gas stream. Inspiratory flows decrease as the air-entrainment port is closed.

PTS:	1	REF:	Page 164

43.	A patient wants to use a spinning disk nebulizer at home. The patient should be aware that it:

| a. | increases the ambient relative humidity. |
| --- | --- |
| b. | is a mechanical aerosol generator. |
| c. | can be a source of bacterial contamination to the environment. |
| d. | can deliver 20 mL/min. |

ANS:	C

The room nebulizers are not commonly recommended because they can be a source of bacterial contamination to the environment. See also Figure 6-11.

PTS:	1	REF:	Page 164

44.	The respiratory therapist needs to deliver a bland aerosol to a 24-month-old child through a spinning disk device. The appropriate flow setting is ___ L/min to ___ L/min.

| a. | 3; 5 | c. | 10; 15 |
| --- | --- | --- | --- |
| b. | 6; 8 | d. | 20; 22 |

ANS:	D

Larger versions of these devices (with 2- to 3-L reservoirs) are used to deliver bland aerosols into mist tents. These enclosure systems can generate flow rates greater than 20 L/min, with water outputs as high as 5 mL/min (300 mL/hr). Because heat buildup in enclosures is a problem, these systems are always unheated.

PTS:	1	REF:	Pages 164-165

45.	A sputum induction is needed for a 55-year-old female patient. Which of the following devices will be most appropriate?

| a. | Spinning disk nebulizer | c. | Ultrasonic nebulizer |
| --- | --- | --- | --- |
| b. | Wick humidifier | d. | Large-volume jet nebulizer |

ANS:	C

Ultrasonic nebulizers are valued clinical tools for inducing sputum production so that specimens can be used for diagnostic analysis. Ultrasonic nebulizers yield a higher quantity and quality of sputum specimens for analysis than other nebulizers but at some cost of increased airway reactivity.

PTS:	1	REF:	Page 166

46.	The device that presents the least risk of cross-contamination as a room humidifier is which of the following?

| a. | Bubble humidifier | c. | Wick humidifier |
| --- | --- | --- | --- |
| b. | Large-volume nebulizer | d. | Ultrasonic nebulizer |

ANS:	C

Passover and wick type of humidifiers present less risk of cross-contamination than does the ultrasonic nebulizer as a room humidifier, which can easily become contaminated, resulting in airborne transmission of pathogens.

PTS:	1	REF:	Page 166

47.	You have to set up a continuous heated aerosol to a patient with a tracheostomy. Which of the following pieces of equipment will be needed?

1. Aerosol mask

2. Tracheostomy collar

3. Thermometer

4. Drain bag

5. Oxygen tubing

6. Corrugated tubing

7. Large-volume nebulizer

8. Ultrasonic nebulizer

9. Heating unit

10. Bubble humidifier

| a. | 1, 5, and 10 | c. | 1, 4, 6, 7, and 9 |
| --- | --- | --- | --- |
| b. | 2, 4, 6, and 8 | d. | 2, 3, 4, 6, 7, and 9 |

ANS:	D

The appropriate pieces of equipment needed to deliver a continuous heated aerosol are listed in Box 6-5. An ultrasonic nebulizer should not be used because it has an increased risk of causing bronchospasm with continuous use.

PTS:	1	REF:	Page 167, Box 6-5

48.	A patient has a heated air-entrainment large-volume jet nebulizer. The aerosol disappears completely during each inspiration. The most-appropriate action to take is to:

| a. | decrease the oxygen flow to the large-volume nebulizer. |
| --- | --- |
| b. | increase the oxygen flow to the large-volume nebulizer. |
| c. | decrease the heat level. |
| d. | remove the corrugated reservoir tubing. |

ANS:	B

The aerosol should not completely disappear during inspiration. This means that the device is not meeting the needs of the patient. The oxygen flow can be increased without worry that the fractional inspired oxygen will increase, because the device uses air entrainment to set the oxygen concentration. This is the only answer that will increase the output of the device to meet the patient’s inspiratory needs.

PTS:	1	REF:	Page 167, Box 6-5

49.	You have been called because of an audible alarm from a heated air-entrainment large-volume jet nebulizer. The most-appropriate immediate action is to:

| a. | decrease the oxygen flow. |
| --- | --- |
| b. | change the device to an injection nebulizer. |
| c. | change the device to a high-flow heated humidifier. |
| d. | turn off the heater power to prevent overheating. |

ANS:	D

Audible and visual alarms should indicate when remote temperature sensors are disconnected, absent, or defective; in these cases, power to the heater should be interrupted to prevent overheating. See Box 6-7.

PTS:	1	REF:	Page 176, Box 6-7

50.	A heated air-entrainment large-volume jet nebulizer needs to be set up for a patient. In addition to the nebulizer, which of the following pieces of equipment are needed?

1. Heating device

2. Tracheostomy collar

3. Thermometer

4. Aerosol mask

5. Drain bag

6. Oxygen tubing

7. Corrugated tubing

| a. | 1, 2, 5, and 7 | c. | 3, 4, 5, and 6 |
| --- | --- | --- | --- |
| b. | 4 and 7 | d. | 1, 3, 4, 5, and 7 |

ANS:	D

Because the patient is not intubated, an aerosol mask needs to be used with a nebulizer that has a heating device. Corrugated tubing with a thermometer and a drain bag in-line to collect the rain out are also needed. Oxygen tubing is not used with large-volume nebulizers.

PTS:	1	REF:	Page 167, Box 6-5

51.	Which of the following is not a complication from bland aerosol therapy?

| a. | Infection | c. | Bronchospasm |
| --- | --- | --- | --- |
| b. | Overhydration | d. | Upper airway edema |

ANS:	D

The most common problems with bland aerosol delivery systems involve infection control, environmental safety, inadequate mist production, overhydration, bronchospasm, and noise.

PTS:	1	REF:	Page 200, Clinical Practice Guideline 6-2

52.	Which of the following presents the greatest risk for an infant being treated with a heated jet nebulizer?

| a. | Infection | c. | Overhydration |
| --- | --- | --- | --- |
| b. | Bronchospasm | d. | Tachycardia |

ANS:	C

Overhydration is a potentially serious problem for patients being treated with heated jet nebulizers and ultrasonic nebulizers. The risk of overhydration is greatest for infants, young children, and those with preexisting fluid or electrolyte imbalances.

PTS:	1	REF:	Page 169

53.	An ultrasonic nebulizer is being used to deliver a hypotonic solution. The patient suddenly reports shortness of breath. The respiratory therapist auscultates, which reveals wheezing over both lung fields. Which of the following actions should be taken first?

1. Observe the patient.

2. Give the patient oxygen.

3. Add a bronchodilator to the treatment.

4. Immediately stop the treatment.

| a. | 1, 2, and 3 | c. | 2 and 4 |
| --- | --- | --- | --- |
| b. | 1 and 3 | d. | 4 |

ANS:	D

If bronchospasm occurs during therapy, treatment must be stopped immediately, oxygen must be provided, and appropriate bronchodilator therapy should be initiated as soon as possible.

PTS:	1	REF:	Page 169

54.	A hypotonic solution with an ultrasonic nebulizer is ordered for a patient who has a history of hyperreactive airways. In such a situation, the respiratory therapist should recommend which of the following?

1. Refuse to give the treatment.

2. Suggest giving a bronchodilator before the hypotonic solution.

3. Suggest that an isotonic solution be used instead.

4. Administer oxygen before the hypotonic solution.

| a. | 1 | c. | 3 and 4 |
| --- | --- | --- | --- |
| b. | 2 and 3 | d. | 4 |

ANS:	B

If the physician still requests bland aerosol therapy for this patient, pretreatment with a bronchodilator might be required. In addition, isotonic solutions (0.9% saline) might also be better tolerated by these patients than water. Administered oxygen will not prevent bronchospasm—even if an isotonic solution is used in a patient with hyperreactive airways.

PTS:	1	REF:	Pages 169-170

55.	An individual’s humidity deficit will be the smallest when breathing:

| a. | oxygen through a heated-wick humidifier. |
| --- | --- |
| b. | room air through a heat and moisture exchanger. |
| c. | room air at 30 C and 50% relative humidity. |
| d. | by nasal cannula (6 L/min) with a bubble humidifier. |

ANS:	A

A wick humidifier can deliver 44 mg/L of water vapor as compared with 10 to 31 mg/L from a heat and moisture exchanger, 15 mg/L from room air at 30 C and 50% relative humidity, and 10 to 20 mg/L from a bubble humidifier.

PTS:	1	REF:	Page 159, Table 6-1

56.	Which of the following devices would be the least effective in reducing a patient’s humidity deficit while intubated?

| a. | A heat and moisture exchanger | c. | A wick humidifier set at 35° C |
| --- | --- | --- | --- |
| b. | An unheated bubble humidifier | d. | Heated air-entrainment nebulizer |

ANS:	B

The unheated bubble humidifier has a water vapor output of 10 to 20 mg H2O/L compared with the heat and moisture exchanger, with approximately 30 mg H2O/L; the wick humidifier, with 44 mg H2O/L; and the heated air-entrainment nebulizer, with 33 to 55 mg H2O/L.

PTS:	1	REF:	Page 167

57.	Which of the following is true concerning heat and moisture exchangers?

| a. | They are passover type A humidifiers. |
| --- | --- |
| b. | They might reduce the incidence of nosocomial infections. |
| c. | They are best-suited for long-term use on mechanical ventilators. |
| d. | They are able to provide a humidity output of 10 mg/L to 44 mg/L at 35 C. |

ANS:	B

Heat and moisture exchangers might reduce the incidence of nosocomial infections. Passive humidifier heat and moisture exchangers have been successfully used to meet the short-term humidification needs of spontaneously breathing and mechanically ventilated patients with endotracheal and tracheostomy tubes. However, a heat and moisture exchanger is not a passover humidifier. The passover humidifier directs gas over a liquid or liquid-saturated surface. Heat and moisture exchanging filters have an increased internal surface area and are capable of providing moisture output in the range of 18 mg H2O/L to 28 mg H2O/L. According to the American Association for Respiratory Care’s Clinical Practice Guideline, heat and moisture exchangers should not be used for long-term humidification of patients who are receiving invasive mechanical ventilation.

PTS:	1	REF:	Page 175

58.	Creation of aerosol particles within a small volume nebulizer is based on:

| a. | Stokes’ law. | c. | the Bernoulli principle. |
| --- | --- | --- | --- |
| b. | Brownian motion. | d. | Reynolds’ number. |

ANS:	C

A typical small-volume nebulizer is powered by a high-pressure stream of gas directed through a restricted orifice (the jet). The gas stream leaving the jet passes the opening of a capillary tube immersed in solution. Because it produces low lateral pressure at the outlet, the high jet velocity draws the liquid up the capillary tube and into the gas stream, where it is sheared and broken up into droplets. This is known as the Bernoulli principle.

PTS:	1	REF:	Page 192

59.	Driving a small-volume nebulizer with heliox, instead of air or oxygen, has what effect on the aerosol produced?

| a. | Increases upper airway impaction |
| --- | --- |
| b. | Creates larger particles |
| c. | Decreases the aerosol output |
| d. | Increases the turbulence of the aerosol |

ANS:	A

The density of the driving gas affects both aerosol generation and the delivery of aerosols to the lungs. Lowering the density of the carrier gas reduces the amount of turbulent flow, which results in less aerosol impaction in the upper airways. When heliox is used to drive a jet nebulizer at standard flow rates, aerosol output is substantially less than with air or oxygen, and aerosol particles are considerably smaller.

PTS:	1	REF:	Page 193

60.	To produce clinically useful aerosol particles through a small-volume nebulizer, the gas flow rate should be set at ___ L/min to ___ L/min.

| a. | 12; 15 | c. | 6; 10 |
| --- | --- | --- | --- |
| b. | 10; 12 | d. | 3; 6 |

ANS:	C

Set flow to the manufacturer’s recommendations; often this is 6 to 10 L/min. See Box 4-11.

PTS:	1	REF:	Page 198, Box 6-11

61.	Which of the following is the optimal technique for use of a metered dose inhaler?

1. Actuating the inhaler with the mouthpiece 4 cm from the patient’s opened mouth

2. Placing the inhaler in the mouth and closing the lips tightly

3. Holding breath for 5 to 10 seconds after inhalation

4. Not waiting between actuations

5. Inhaling the aerosol quickly

| a. | 1 and 3 | c. | 2, 3, and 5 |
| --- | --- | --- | --- |
| b. | 1, 3, and 4 | d. | 2, 4, and 5 |

ANS:	A

If an adult patient positions the outlet of the pressurized metered dose inhaler approximately 4 cm (two finger widths) in front of an opened mouth and has a low inspiratory flow rate, this technique can increase the dose delivered to the lower respiratory tract from a range of approximately 7% to 10% to 14% to 20% and can decrease the impact on the oropharyngeal region.

PTS:	1	REF:	Page 186, Box 6-9

62.	Which of the following are true about the use of a metered dose inhaler?

1. Inertial impaction can be reduced if the patient holds the inhaler in the mouth.

2. Smaller particles will be produced in a cold environment.

3. Larger particles will be produced in a cold environment.

4. The optimum temperature for an inhaler canister is 37° C.

5. The inhaler needs to be shaken before actuation.

| a. | 1, 3, and 4 | c. | 2 and 5 |
| --- | --- | --- | --- |
| b. | 2 and 4 | d. | 3, 4, and 5 |

ANS:	D

If an adult patient positions the outlet of the pressure metered dose inhaler approximately 4 cm (two finger widths) in front of an opened mouth with a low inspiratory flow rate, this practice can increase the dose delivered to the lower respiratory tract from a range of approximately 7% to 10% to 14% to 20% and decrease oropharyngeal impaction. A cold ambient temperature (i.e., a temperature of less than 10 C) dramatically decreases the output of chlorofluorocarbon pressure metered dose inhalers.

PTS:	1	REF:	Page 186, Box 6-9

63.	One criterion for the use of a dry-powder inhaler is the patient’s ability to:

| a. | coordinate inspiration with actuation. |
| --- | --- |
| b. | hold the dry-powder inhaler 4 cm from the mouth. |
| c. | empty the powder into the dry-powder inhaler. |
| d. | inspire at flows greater than 40 L/min. |

ANS:	D

The most critical factor in passive dry-powder inhaler use is the need for high inspiratory flow. Patients must generate inspiratory flows of at least 40 to 60 L/min to produce a respirable powder aerosol.

PTS:	1	REF:	Page 188

64.	Which of the following devices can be used to administer a continuous bland aerosol?

1. Face tent

2. Simple mask

3. Nonrebreathing mask

4. Tracheostomy collar

| a. | 1 | c. | 2 and 3 |
| --- | --- | --- | --- |
| b. | 1 and 4 | d. | 2, 3, and 4 |

ANS:	B

The face tent and tracheostomy collar are patient interface appliances used to deliver bland aerosol therapy. The others are not.

PTS:	1	REF:	Page 167

65.	A patient with subglottic edema might benefit from what type of humidity or aerosol therapy?

| a. | Continuous heated bland aerosol |
| --- | --- |
| b. | Humidity from a bubble humidifier |
| c. | Continuous cold-nebulized bland aerosol |
| d. | Dry-powder inhaler with a bronchodilator |

ANS:	C

A continuous cold-nebulized bland aerosol will produce vasoconstriction and reduce edema. Continuous heated bland aerosol is used for patients with hypothermia. A dry-powder inhaler with a bronchodilator is used to relieve a bronchospasm. Humidity from a bubble humidifier will not reduce subglottic edema.

PTS:	1	REF:	Page 199, Clinical Practice Guideline 6-1

66.	Which of the following offers the best method for delivering ribavirin?

| a. | A small-volume nebulizer and an aerosol mask |
| --- | --- |
| b. | An LVN, large-bore tubing, and a Briggs adapter |
| c. | SPAG-2, large-bore tubing, and an oxygen hood |
| d. | SPAG-2, large-bore tubing, and an aerosol mask |

ANS:	C

The SPAG was manufactured by ICN Pharmaceuticals specifically for the administration of ribavirin (Virazole) to infants with respiratory syncytial virus infection.

PTS:	1	REF:	Page 195

67.	Continuous bland aerosol administration can be hazardous when administered to patients with:

1. asthma.

2. pneumonia.

3. cystic fibrosis.

4. ineffective cough mechanisms.

| a. | 1 and 3 | c. | 2 and 4 |
| --- | --- | --- | --- |
| b. | 1 and 4 | d. | 2, 3, and 4 |

ANS:	B

Continuous bland aerosols should not be used for patients with asthma because of bronchospasm. For patients who have ineffective cough mechanisms, continuous aerosol use can cause overhydration.

PTS:	1	REF:	Page 200, Clinical Practice Guideline 6-2

68.	Continuous bland aerosol is being delivered to an intubated patient through a large-volume nebulizer and a Briggs adapter. A clinical assessment reveals bronchoconstriction. Which of the following is the most appropriate recommendation?

| a. | Switch to an ultrasonic nebulizer. |
| --- | --- |
| b. | Change to a heated-wick humidifier. |
| c. | Add a heater to the large-volume nebulizer. |
| d. | Decrease the flow rate to the large-volume nebulizer. |

ANS:	B

Change to a heated-wick humidifier, which is a heated humidifier that is appropriate when airways are hyperreactive. Ultrasonic nebulizers and large-volume nebulizers will worsen bronchospasms because they also create bland aerosol.

PTS:	1	REF:	Page 178, Figure 6-28

69.	A metered dose inhaler should be held away from the mouth during actuation to:

| a. | decrease the effect of inertial impaction on the particles. |
| --- | --- |
| b. | decrease the effect of gravity on the particles. |
| c. | ensure deposition in the oropharynx. |
| d. | prevent cross-contamination. |

ANS:	A

This technique will decrease the effect of inertial impaction on the particles in the oropharynx.

PTS:	1	REF:	Page 183

70.	Which of the following will correct the problem of tubing condensation within the ventilator circuit?

| a. | Water traps | c. | Hot-wire circuit |
| --- | --- | --- | --- |
| b. | Increase the room temperature | d. | Both A and C |

ANS:	D

Hot-wire circuits are quite common in the intensive care unit. The placement of wire-heating elements into the ventilator circuit allows the gas delivered to the patient to remain at a constant temperature.

PTS:	1	REF:	Page 177

71.	Which of the following are contraindications for heat moisture exchanger (HME) use?

1. Hypothermia

2. Uncuffed endotracheal tubes

3. With heated humidification

4. Decrease volume of secretions

| a. | 1, 2, and 3 | c. | 2 and 3 |
| --- | --- | --- | --- |
| b. | 1 and 3 | d. | 2 and 4 |

ANS:	A

Contraindications of heat and moisture exchanger use are hypothermia, the presence of uncuffed endotracheal tubes, use with heated humidification, and an increase in the volume of secretions.

PTS:	1	REF:	Page 164, Figure 6-7

72.	Which of the following are valid reasons for not using a heat and moisture exchanger in infants?

1. Heat and moisture exchangers add 30 mL to 90 mL of mechanical dead space.

2. Heat and moisture exchangers are not effective with uncuffed endotracheal tubes.

3. Heat and moisture exchangers might cause an obstruction in endotracheal tubes in infants.

4. Heat and moisture exchangers are an inexpensive alternative to humidifiers.

| a. | 1, 2, and 3 | c. | 2 and 4 |
| --- | --- | --- | --- |
| b. | 1 and 4 | d. | 3 and 4 |

ANS:	A

Heat and moisture exchangers are not recommended for use in infants for several reasons. Heat and moisture exchangers add 30 to 90 mL of mechanical dead space, often exceeding the tidal volume of a 5-kg infant (i.e., 25 mL). Uncuffed endotracheal tubes used in infants allow a portion of exhaled gas to leak around the tube, bypassing the heat and moisture exchanger and reducing its ability to capture exhaled heat and humidity. Providing heated humidity at less than 35 C has been associated with narrowing and obstruction of endotracheal tubes in infants. The advantages of heat and moisture exchangers are that they are inexpensive alternatives to humidifiers and they are the best choice for adult patients.

PTS:	1	REF:	Page 176

73.	A heat and moisture exchanger is best placed where in the ventilator circuit?

| a. | Directly at the patient Y-connector |
| --- | --- |
| b. | Twenty centimeters away from the endotracheal tube |
| c. | At the ventilator outlet |
| d. | At the ventilator inlet |

ANS:	A

Heat and moisture exchangers should be placed directly at the airway or patient Y-connector. Other heat and moisture exchanger locations that might also prove effective are 10 cm away from the endotracheal tube.

PTS:	1	REF:	Page 175

74.	Use of a heat and moisture exchanger is beneficial to a patient with a(n):

| a. | tracheostomy who is receiving low-flow oxygen. |
| --- | --- |
| b. | tracheostomy who is receiving room air. |
| c. | open stoma. |
| d. | All of the above. |

ANS:	D

Patients who no longer need an artificial airway—but who breathe through an open stoma—might benefit from heat and moisture exchangers. Heat and moisture exchangers have been developed for use in patients with tracheostomies who no longer require mechanical ventilation; these patients use only room air or low-flow oxygen.

PTS:	1	REF:	Page 175

75.	Which of the following offers the most accurate and reliable method for measuring the temperature and humidity level within a ventilator circuit?

| a. | Humidity control of hot-wire humidification system |
| --- | --- |
| b. | Oxygen analyzer |
| c. | Hygrometer-thermometer |
| d. | Not able to measure humidity level |

ANS:	C

The most accurate and reliable way to ensure that patients are receiving gas at the expected temperature and humidity level is to measure these parameters with the portable battery-operated digital hygrometer-thermometer system. Humidity control on a hot-wire humidification system will not provide a measurement of absolute humidity or relative humidity; instead, it will enable the measurement of only the temperature differential between the humidifier and the airway sensor.

PTS:	1	REF:	Page 177

76.	It is appropriate to suggest which of the following to provide humidification for an intubated patient with a cuffed endotracheal tube, no secretions, and a tidal volume of 500 mL?

| a. | A heat and moisture exchanger | c. | Heated aerosol with humidifier |
| --- | --- | --- | --- |
| b. | Heated aerosol or humidifier | d. | Systemic hydration |

ANS:	A

If there are no secretions, an HME would be suitable.    Use the algorithm for selecting the correct humidity and bland aerosol therapy located in the text.

PTS:	1	REF:	Page 178, Figure 6-28

77.	Which of the following devices or methods would be appropriate to provide sufficient humidity for an intubated patient with an uncuffed endotracheal tube who has thick, bloody secretions?

| a. | Heat and moisture exchanger |
| --- | --- |
| b. | Heated humidifier |
| c. | Heated humidifier with a heat and moisture exchanger |
| d. | Systemic hydration |

ANS:	B

According to the algorithm for selecting the correct humidity and bland aerosol therapy, a heated humidifier is appropriate for this patient. Both the use of an uncuffed endotracheal tube and the presence of thick, bloody secretions are contraindications for the use of a heat and moisture exchanger. Heat and moisture exchangers cannot be used with heated humidification. Systemic hydration does not adequately humidify inhaled gas.

PTS:	1	REF:	Page 176, Box 6-7

78.	Which of the following treatments would provide sufficient humidity to a conscious patient with a core temperature less than 36° C who had nearly drowned?

| a. | Heat and moisture exchanger | c. | Heated aerosol with humidifier |
| --- | --- | --- | --- |
| b. | Heated aerosol or humidifier | d. | Systemic hydration |

ANS:	B

Because the patient is hypothermic, a heated aerosol or heated humidifier needs to be used. Hypothermia is a contraindication for the use of a heat and moisture exchanger. Heated aerosol generators are not used in conjunction with humidifiers, and systemic hydration will not help warm the patient’s core temperature.

PTS:	1	REF:	Page 178, Figure 6-28

79.	Which of the following are the gram-negative bacilli that are responsible for nosocomial infections when aerosol generators are in use?

| a. | Pseudomonas aeruginosa and Streptococcus pneumoniae |
| --- | --- |
| b. | Legionella pneumophila and Streptococcus pneumoniae |
| c. | Pseudomonas aeruginosa and Legionella pneumophila |
| d. | Legionella pneumophila and Streptococcus viridans |

ANS:	C

The gram-negative bacilli are Pseudomonas aeruginosa and Legionella pneumophila. All the other choices are gram-positive cocci.

PTS:	1	REF:	Page 180

80.	Which of the following is appropriate to provide humidity to a patient with cold air reactive airways?

| a. | HME | c. | Heated humidifier |
| --- | --- | --- | --- |
| b. | Heated aerosol | d. | Systemic hydration |

ANS:	C

A patient with cold air reactive airways would be best suited with a heated humidifier according to the algorithm for selecting the correct humidity and bland aerosol therapy.

PTS:	1	REF:	Page 178, Figure 6-28

81.	Reactive bronchospasm and increased airway resistance can be caused by:

| a. | warm humidity. | c. | aerosolized normal saline. |
| --- | --- | --- | --- |
| b. | warm aerosol. | d. | cold high-density aerosol. |

ANS:	D

The administration of cold high-density aerosols increases the risk of developing reactive bronchospasm and increased airway resistance.

PTS:	1	REF:	Page 180

82.	Which of the following represent the immediate signs and symptoms of a patient experiencing bronchospasm during the administration of a bland aerosol?

1. Percentage of forced expiratory volume in 1 second (%FEV1) <75%

2. Adventitious breath sounds

3. Increased work of breathing

4. Change of breathing pattern

| a. | 1, 2, and 3 | c. | 2, 3, and 4 |
| --- | --- | --- | --- |
| b. | 1, 3, and 4 | d. | 1, 2, 3, and 4 |

ANS:	D

Monitoring for reactive bronchospasm should consist of peak flow measurements or %FEV1 before and after therapy; auscultation for adventitious breath sounds; observation of the patient’s breathing pattern and overall appearance; and, most essentially, communication with the patient during therapy to determine the perceived work of breathing.

PTS:	1	REF:	Page 180

83.	Which of the following statements about breath-actuated nebulizers is not true?

| a. | They can be used during the entire respiratory cycle. |
| --- | --- |
| b. | They do not require good hand-breath coordination. |
| c. | They decrease aerosol medication waste. |
| d. | They are not appropriate for some patients with severe airway obstruction. |

ANS:	A

Breath-actuated nebulizers generate aerosol only during inspiration. It is true that breath-actuated pressure metered dose inhalers do not require good hand-breath coordination, and because of that they decrease aerosol medication waste. Patients with severe airway obstruction may not have the capability of generating enough flow to trigger the breath actuator.

PTS:	1	REF:	Pages 190-191

84.	Which of the following statements is true about the use of a pressure metered dose inhaler with a valved holding chamber?

| a. | This chamber cannot be used with infants. |
| --- | --- |
| b. | Infants can take up to 10 breaths. |
| c. | Shake the pressure metered dose inhaler and chamber separately. |
| d. | This device is only for adults and children over the age of 12. |

ANS:	B

A pressure metered dose inhaler with a valved holding chamber can be used by adults and infants. Adults will take 1 to 3 breaths, and infants can take up to 10 breaths. Shake the pressure metered dose inhaler with the chamber.

PTS:	1	REF:	Page 183, Box 6-8

85.	Which of the following statements is true concerning the use of dry-powder inhalers?

| a. | Propellants are used to help deliver the drug. |
| --- | --- |
| b. | The carrier substance is usually lactose or glucose. |
| c. | Laminar flow is necessary to create the aerosol. |
| d. | Patients must breathe slowly and deeply during administration. |

ANS:	B

The carrier substance is usually lactose or glucose. Turbulent flow is a function of the ability of the patient to inhale the powder with a sufficiently high inspiratory flow rate. Propellants are not used in dry-powder inhalers.

PTS:	1	REF:	Page 188

86.	A respiratory therapist needs to deliver a bronchodilator with a small-volume nebulizer. To ensure that 70% of the medication is available for nebulization, _____ mL of fluid should fill the nebulizer.

| a. | 3 | c. | 3.5 |
| --- | --- | --- | --- |
| b. | 5 | d. | 4 |

ANS:	B

Filling the nebulizer with 5 mL would make 3.5 mL, or 70% of the medication, available for nebulization.

PTS:	1	REF:	Page 192

87.	Which of the following statements is true concerning vibrating mesh nebulizers?

| a. | They require at least 3 mL of solution to produce an aerosol. |
| --- | --- |
| b. | The exit velocity of the aerosol from vibrating mesh nebulizers is rapid. |
| c. | They add no gas to the aerosol stream. |
| d. | They can produce a particle size range of 1 to 2 m. |

ANS:	C

Vibrating mesh nebulizers add no gas to the aerosol stream. They can generate aerosol from a single drop of liquid, so they require very little solution to produce an aerosol. The exit velocity of the aerosol is slow (less than 4 m/s), and the particle can range from 2 to 5 m (mass median aerodynamic diameter).

PTS:	1	REF:	Page 193

88.	Which of the following statements about vibrating mesh nebulizers is false?

| a. | They have no residual volume. |
| --- | --- |
| b. | They reduce the contamination of medication in the reservoir. |
| c. | They are capable of generating aerosol from a single drop of liquid. |
| d. | They can be used in closed systems during mechanical ventilation. |

ANS:	A

Vibrating mesh nebulizers do have residual drug volumes that range from 0.1 to 1.2 mL. As is not the case with jet nebulizers and ultrasonic nebulizers, the mesh acts as a physical barrier between the drug in the medication reservoir and the body of the nebulizer in which the aerosol output is collected. This reduces the risk that exhaled pathogens will contaminate the medication in the reservoir. They can generate an aerosol from a single drop of liquid.

PTS:	1	REF:	Page 193

89.	A patient who requires mechanical ventilation has a core body temperature of 32o C; humidity is being provided with a heat moisture exchanger. Which of the following statements is accurate?

| a. | The relative humidity at BTPS will be increased. |
| --- | --- |
| b. | Higher inspiratory flow rates will be required. |
| c. | Humidification will be inadequate due to the decreased body temperature. |
| d. | The minute volume will be increased due to the hypothermia. |

ANS:	C

Prolonged and severe shifts in the isothermic saturation boundary can damage the airways and compromise the body’s normal heat and moisture exchange mechanisms.

REF Page 6-10

PTS:	1	REF:	Page 161