Chapter 14: Infant and Pediatric Devices

MULTIPLE CHOICE

1.	What is the respiratory rate available on the CareFusion SiPAP?

| a. | 1 to 100 breaths/minute | c. | 1 to 150 breaths/minute |
| --- | --- | --- | --- |
| b. | 1 to 125 breaths/minute | d. | 1 to 200 breaths/minute |

ANS:	C

The respiratory rate on the CareFusion SiPAP is adjustable from 0 to 150 breaths per minute.

PTS:	1	REF:	Page 466, Table 14-1

2.	When is the oxygen blender on the CareFusion Infant Flow Nasal CPAP System recalibrated?

| a. | When the calibration control is activated |
| --- | --- |
| b. | When the ventilator is turned on |
| c. | When there is a change in the delivered fractional inspired oxygen |
| d. | At start-up and every 12 hours |

ANS:	B

The oxygen blender is automatically recalibrated on start-up.

PTS:	1	REF:	Page 465

3.	Which of the following are power sources for the Cardinal Health Infant Flow Nasal CPAP System?

1. 12-V direct current battery

2. 20-psig compressed-air source

3. 50-psig oxygen source

4. Alternating current electrical source

| a. | 1 and 4 | c. | 2, 3, and 4 |
| --- | --- | --- | --- |
| b. | 1 and 2 | d. | 1, 2, 3, and 4 |

ANS:	D

The continuous positive airway pressure system uses two high-pressure gas sources (50 psig each), an alternating current electrical source, and a 12-V direct current battery as a back-up electrical power source.

PTS:	1	REF:	Page 464

4.	The Cardinal Health Infant Flow Nasal CPAP System driver can operate on battery power for _____ hours.

| a. | 2 | c. | 6 |
| --- | --- | --- | --- |
| b. | 4 | d. | 10 |

ANS:	B

A lead acid battery is contained within the driver. With a full charge, the driver unit operates for 4 hours on battery power.

PTS:	1	REF:	Page 464

5.	What is the maximum amount of time an alarm condition must persist on the CareFusion Infant Flow Nasal CPAP System for an alarm to be delayed when the ventilator is being set up?

| a. | 30 seconds | c. | 90 seconds |
| --- | --- | --- | --- |
| b. | 60 seconds | d. | 120 seconds |

ANS:	D

Alarms are delayed for 120 seconds after entering the main screen to allow time for ventilator set up.

PTS:	1	REF:	Page 465

6.	The CareFusion Airlife Nasal CPAP System Driver is able to operate with its built-in battery for approximately _____ hour(s).

| a. | 1 | c. | 3 |
| --- | --- | --- | --- |
| b. | 2 | d. | 4 |

ANS:	D

The built-in battery has an estimated 4 hours of back-up power when alternating current power is unavailable.

PTS:	1	REF:	Page 464

7.	The “comfort-flow feature” on the CareFusion AirLife Nasal CPAP System Driver refers to which of the following?

| a. | Allows humidification to be used with the unit |
| --- | --- |
| b. | Enables the unit to alternate between two levels of pressure and flow |
| c. | Allows gas flow reduction on exhalation to reduce resistance |
| d. | Automatically adjusts gas flow by the unit to achieve the desired pressure |

ANS:	D

This unit automatically adjusts the gas flow to achieve the desired pressure levels and to compensate for leaks. A constant level of nasal continuous positive airway pressure is maintained.

PTS:	1	REF:	Page 465

8.	Which of the following statements is true concerning the CareFusion AirLife Nasal CPAP System Driver?

| a. | The internal oxygen cell analyzes fractional inspired oxygen and automatically calibrates. |
| --- | --- |
| b. | The desired continuous positive airway pressure level is adjusted by moving a probe in 1-cm increments. |
| c. | Bilevel airway pressure is available to ventilate newborns with apnea of prematurity. |
| d. | The proximal pressure line of the delivery circuit attaches to the port on the front panel. |

ANS:	A

The internal oxygen cell analyzes the fractional inspired oxygen and automatically calibrates oxygen binding on start-up. Adjusting a probe by 1-cm increments is done for the Fisher & Paykel Bubble CPAP System. Bilevel is available only on the CareFusion Infant Flow SiPAP unit. The CareFusion Infant Flow Nasal CPAP System has a proximal pressure line.

PTS:	1	REF:	Page 465

9.	Which alarms remain audible in the standby mode on the CareFusion AirLife Nasal CPAP System Driver?

1. Low flow

2. Disconnect

3. High pressure

4. Low oxygen pressure

| a. | 1 and 2 | c. | 2 and 4 |
| --- | --- | --- | --- |
| b. | 1 and 3 | d. | 3 and 4 |

ANS:	B

While in the standby mode, the system will operate normally, with all audible alarms muted except for the high-pressure and low-flow alarms.

PTS:	1	REF:	Page 465

10.	The maximum number of hours that the CareFusion AirLife Nasal CPAP System Driver can display trend data is _____ hours.

| a. | 4 | c. | 12 |
| --- | --- | --- | --- |
| b. | 8 | d. | 24 |

ANS:	D

The unit has a trend screen that can display data for the previous 4, 8, 12, or 24 hours, with 24 hours being the maximum amount of time.

PTS:	1	REF:	Page 465

11.	Which of the following units is able to provide more support than traditional noninvasive continuous positive airway pressure to transition infants and pediatric patients from intubation to successful extubation?

| a. | CareFusion Infant Flow SiPAP System |
| --- | --- |
| b. | CareFusion Infant Flow Nasal CPAP System |
| c. | Fisher & Paykel Healthcare Bubble CPAP System |
| d. | CareFusion Health AirLife Nasal CPAP System Driver |

ANS:	A

The CareFusion Infant Flow SiPAP System can provide noninvasive bilevel airway pressure that can bridge the gap between continuous positive airway pressure and mechanical ventilation.

PTS:	1	REF:	Pages 468-469

12.	Which of the following statements concerning the CareFusion/Infant Flow SiPAP Comprehensive System is not true?

| a. | This unit uses an abdominal transducer to detect apnea. |
| --- | --- |
| b. | This unit is able to provide time-triggered biphasic ventilation. |
| c. | This unit is able to provide patient-triggered biphasic ventilation. |
| d. | The battery can provide 16 hours of operation if alternating current power is lost. |

ANS:	D

The CareFusion Infant Flow SiPAP Comprehensive System can provide only up to 2 hours of operating time if the alternating current power source is lost. The battery requires 16 hours to become fully charged from a discharged state.

PTS:	1	REF:	Pages 465-466

13.	Patient inspiration causes which of the following to occur on the CareFusion Infant Flow SiPAP System?

| a. | The NOR/OR fluidic gate closes. |
| --- | --- |
| b. | Flow is directed toward the expiratory limb. |
| c. | Flow is taken away from the inspiratory limb. |
| d. | Kinetic energy of flow is converted to pressure. |

ANS:	D

When the patient makes an inspiratory effort, the flow generator converts the kinetic energy of flow to pressure to stabilize pressure delivery at the patient interface. The unit uses flip-flop fluidic logic.

PTS:	1	REF:	Page 467, Table 14-7

14.	The CareFusion SiPAP is set to deliver an inspiratory positive airway pressure of 6 cm H2O and an expiratory positive airway pressure of 3 cm H2O. The high airway pressure alarm would activate at which of the following pressures?

| a. | 9 cm H2O | c. | 17 cm H2O |
| --- | --- | --- | --- |
| b. | 11 cm H2O | d. | 21 cm H2O |

ANS:	A

The high airway pressure alarm activates with a pressure increase 3 cm H2O over the set pressure.

PTS:	1	REF:	Page 468, Table 14-2

15.	A low battery charge alarm condition on a CareFusion SiPAP is caused by which of the following?

| a. | Battery charge is totally drained. | c. | Battery charge fell below 40%. |
| --- | --- | --- | --- |
| b. | Battery charge fell below 75%. | d. | Battery voltage fell below 10 V. |

ANS:	C

The low battery charge alarm activates if battery charge falls below 40%.

PTS:	1	REF:	Page 468, Table 14-2

16.	The range of continuous positive airway pressure that the Fisher & Paykel Healthcare Bubble CPAP System can provide is _____ cm H2O.

| a. | 1-10 | c. | 3-10 |
| --- | --- | --- | --- |
| b. | 2-8 | d. | 5-15 |

ANS:	C

The Fisher & Paykel Healthcare Bubble CPAP System has a continuous positive airway pressure range of 3-10 cm H2O.

PTS:	1	REF:	Page 469

17.	The Fisher & Paykel Healthcare Bubble CPAP System requires which of the following power sources?

1. 110-V alternating current main power

2. 12-V direct current battery

3. Air/oxygen blender

4. Air and oxygen high-pressure gas sources

| a. | 3 | c. | 1 and 2 |
| --- | --- | --- | --- |
| b. | 1 and 4 | d. | 2, 3, and 4 |

ANS:	A

The Fisher & Paykel Healthcare Bubble CPAP System requires an air/oxygen blender to power the unit. Only the humidifier requires the electrical source.

PTS:	1	REF:	Page 469

18.	The recommended flow rate range for the Fisher & Paykel Healthcare Bubble CPAP System is _____ L/min.

| a. | 2-4 | c. | 6-8 |
| --- | --- | --- | --- |
| b. | 4-6 | d. | 8-10 |

ANS:	C

The manufacturer-recommended flow rate is 6-8 L/min.

PTS:	1	REF:	Page 471

19.	Which of the following statements concerning the use of the Fisher & Paykel Healthcare Bubble CPAP System is true?

| a. | The bubbling in the continuous positive airway pressure generator should be vigorous. |
| --- | --- |
| b. | Blanching around the nostril rim suggests that the prongs are too long. |
| c. | If the prongs are too small, there is a decrease in airway pressure. |
| d. | Place the prongs, with the curved side up, into the infant’s nares. |

ANS:	B

Large prongs decrease blood flow to the area, causing blanching of the rim of the nostrils. The bubbling in the continuous positive airway pressure generator should be gentle and continuous. If the prongs are too small, there will be an increase in airway pressure, and gas will leak from the system. The prongs should be placed, with the curved side down, into the infant’s nostrils.

PTS:	1	REF:	Page 472, Box 14-3

20.	Which of the following is not true of the Hamilton ARABELLA System?

| a. | An internal battery is not available. |
| --- | --- |
| b. | The flow generator is identical to that of the Infant Flow Nasal CPAP System. |
| c. | The set up and operation are exactly the same as the Infant Flow Nasal CPAP System. |
| d. | A higher work of breathing is needed for the Infant Flow Nasal CPAP System. |

ANS:	B

The major difference between the ARABELLA and the CareFusion Infant Flow Nasal CPAP Systems lies in the design of their flow generators.

PTS:	1	REF:	Page 472, Box 14-3

21.	The stable jet flow from the ARABELLA System causes which of the following?

| a. | A decrease in the patient’s work of breathing |
| --- | --- |
| b. | A higher continuous positive airway pressure level for less flow through the system |
| c. | More patient flow to switch from expiration to inspiration |
| d. | More patient expiratory flow to cause a switch to expiration |

ANS:	D

A more stable jet flow seems to require a more substantial patient expiratory flow to cause the device to switch into expiration. The additional flow that is required for the ARABELLA suggests that a patient must work harder to breathe when this system is used.

PTS:	1	REF:	Page 473

22.	What is the maximum available tidal volume for a time-triggered, pressure-limited, time-cycled infant ventilator with the flow rate set at 8 L/min and the inspiratory time at 0.4 second?

| a. | 19 mL | c. | 53 mL |
| --- | --- | --- | --- |
| b. | 22 mL | d. | 88 mL |

ANS:	C

Tidal volume (liters) = [Inspiratory time (seconds)  Flow (L/min)] ÷ 60 = [0.4  8] ÷ 60 = 0.053 L, or 53 mL.

PTS:	1	REF:	Page 474, Box 14-4

23.	What is the maximum tidal volume available for a time-triggered, pressure-limited, time-cycled infant ventilator with the flow rate set at 6 L/min and the inspiratory time at 0.6 second?

| a. | 13 mL | c. | 53 mL |
| --- | --- | --- | --- |
| b. | 36 mL | d. | 60 mL |

ANS:	D

Tidal volume (liters) = [Inspiratory time (seconds)  Flow (L/min)] ÷ 60 = [0.6  6] ÷ 60 = 0.6 L, or 60 mL.

PTS:	1	REF:	Page 474, Box 14-4

24.	The volume delivered by a time-triggered, pressure-limited, time-cycled infant ventilator depends directly on which of the following ventilator parameters?

1. Flow rate

2. Pressure limit

3. Respiratory rate

4. Inspiratory time

| a. | 1 and 2 | c. | 1, 2, and 4 |
| --- | --- | --- | --- |
| b. | 2 and 3 | d. | 2, 3, and 4 |

ANS:	C

The volume delivered by a pressure-limited, time-triggered, time-cycled ventilator depends on the pressure setting, the flow rate, and the inspiratory time. The respiratory rate is not directly related to volume. It determines the total cycle time and expiratory time because inspiratory time is set. In the case of a set ratio of inspiratory time to expiratory time, the respiratory rate determines the inspiratory time, which is directly related to the volume delivered.

PTS:	1	REF:	Page 474, Box 14-4

25.	The volume delivered by a time-triggered, pressure-limited, time-cycled infant ventilator is dependent on which of the following factors?

1. Positive end-expiratory pressure level

2. Airway resistance

3. Lung compliance

4. Chest wall compliance

| a. | 1 and 3 | c. | 2, 3, and 4 |
| --- | --- | --- | --- |
| b. | 1, 2, and 4 | d. | 1, 2, 3, and 4 |

ANS:	C

Changes in the patient’s airway resistance, lung compliance, and chest wall compliance can alter the volume delivered to the patient. Lower volumes are delivered when airway resistance is high or lung and chest wall compliance are low.

PTS:	1	REF:	Page 474

26.	What is the maximum tidal volume available for a time-triggered, pressure-limited, time-cycled infant ventilator with a set inspiratory time of 0.5 second and a set flow rate of 12 L/min?

| a. | 36 mL | c. | 240 mL |
| --- | --- | --- | --- |
| b. | 100 mL | d. | 250 mL |

ANS:	B

Tidal volume = [Inspiratory time (seconds)  Flow (L/min)] ÷ 60 = [0.5  12 L/min] ÷ 60 = 0.1 L, or 100 mL.

PTS:	1	REF:	Page 474, Box 14-4

27.	Circuit positive end-expiratory pressure may be avoided by using:

| a. | demand flow. | c. | flow interrupter. |
| --- | --- | --- | --- |
| b. | continuous flow. | d. | overpressure-relief valve. |

ANS:	A

Demand flow is believed by some clinicians to be advantageous to operator-selected, continuous-flow systems because continuous-flow systems tend to produce resistance to expiration at the airway, commonly known as circuit positive end-expiratory pressure. For patients whose ventilatory needs include high inspiratory flows but low end-expiratory pressure, the demand system eliminates the need to set a high continuous flow.

PTS:	1	REF:	Page 474

28.	Which of the following mechanisms is incorporated into the CareFusion V.I.P. Bird Infant-Pediatric Ventilator to prevent the risk of auto-PEEP during ventilation with the time-cycled modes?

| a. | Jet Venturi in the exhalation manifold |
| --- | --- |
| b. | Accumulator before the pneumatic regulator |
| c. | Differential pressure transducer by the control diaphragm |
| d. | Pulsation dampener between the regulator and the flow control |

ANS:	A

Because of the possibility of inadvertent positive end-expiratory pressure developing from the continuous flow present in the expiratory limb of the patient circuit during time-cycled, pressure-limited ventilation, a jet Venturi is incorporated into the exhalation manifold.

PTS:	1	REF:	Page 476

29.	What are the two trigger mechanisms for the assist/control volume-cycled mode on the CareFusion V.I.P. Bird Infant-Pediatric Ventilator?

1. Flow

2. Time

3. Volume

4. Pressure

| a. | 1 and 2 | c. | 2 and 4 |
| --- | --- | --- | --- |
| b. | 1 and 3 | d. | 3 and 4 |

ANS:	C

In the assist/control volume-cycled mode, inspiration is time-triggered or pressure-triggered as well as volume-targeted, flow-limited, and volume-cycled.

PTS:	1	REF:	Page 478

30.	With a CareFusion V.I.P. Bird Infant-Pediatric Ventilator, a flashing pressure-support display during spontaneous breathing can be caused by which of the following?

| a. | Pressure support level too high | c. | Leaks around the artificial airway |
| --- | --- | --- | --- |
| b. | High-pressure alarm threshold met | d. | Decreased volume delivery |

ANS:	C

Leaks in the system, including around the artificial airway, may result in excessive inspiratory time. This would cause the ventilator to time-cycle instead of flow-cycle the pressure-support breaths.

PTS:	1	REF:	Page 478

31.	What is the breath termination for a CareFusion V.I.P. Bird Infant-Pediatric Ventilator set to deliver pressure support at tidal volume = 250 mL and peak expiratory flow 15 L/min?

| a. | Time cycle at 1 second | c. | Flow cycle at 1 L/min |
| --- | --- | --- | --- |
| b. | Time cycle at 1.25 second | d. | Flow cycle at 3.75 L/min |

ANS:	D

Breaths in the volume-cycled mode on the CareFusion V.I.P. Bird Infant-Pediatric Ventilator are terminated using flow-cycling. For delivered tidal volume greater than 200 mL, the flow-cycle value is 25% of peak flow.

PTS:	1	REF:	Page 478

32.	The termination sensitivity control on the Care Fusion V.I.P. Bird Infant-Pediatric Ventilator does which of the following?

| a. | Makes it easier to trigger a breath |
| --- | --- |
| b. | Sets the time cycle in the time-cycled modes |
| c. | Sets the flow cycle in the time-cycled modes |
| d. | Adjusts breath termination in pressure support |

ANS:	C

The termination sensitivity control adjusts the flow-termination point of the breath, preventing air-trapping and an inverse inspiratory-to-expiratory ratio, providing expiratory synchrony. It is used only in the assist/control time-cycled mode and has a range of 5% to 25%.

PTS:	1	REF:	Page 478

33.	What is the amount of pressure drop required to initiate the leak-compensation feature on the CareFusion V.I.P. Bird Infant-Pediatric Ventilator during the volume-cycled mode?

| a. | 0.25 cm H2O | c. | 1.00 cm H2O |
| --- | --- | --- | --- |
| b. | 0.50 cm H2O | d. | 1.50 cm H2O |

ANS:	A

When pressure decreases 0.25 cm H2O below baseline pressure, the leak-compensation feature introduces small amounts of flow into the circuit in an attempt to reestablish baseline pressure.

PTS:	1	REF:	Page 479

34.	If a pressure-support breath on a CareFusion V.I.P. Bird Infant-Pediatric Ventilator is 25 mL and the peak inspiratory flow is 10 L/min, the breath will flow-cycle at _____ L/min.

| a. | 0.50 | c. | 1.25 |
| --- | --- | --- | --- |
| b. | 1.00 | d. | 2.00 |

ANS:	A

Between tidal volumes of 0 mL and 50 mL, the flow-cycle value is 5% of peak flow. Therefore, 5%  10 L/min = 0.5 L/min.

PTS:	1	REF:	Page 478, Box 14-6

35.	A pressure-support breath on a CareFusion V.I.P. Bird delivers 230 mL, and the peak inspiratory flow is 18 L/min. This breath will flow-cycle at _____ L/min.

| a. | 0.9 | c. | 3.6 |
| --- | --- | --- | --- |
| b. | 1.8 | d. | 4.5 |

ANS:	D

For delivered volumes greater than 200 mL, the flow cycle value is 25% of peak flow. Therefore, 25%  18 L/min = 4.5 L/min.

PTS:	1	REF:	Page 478, Box 14-6

36.	The CareFusion Bird Partner IIi infant sensor should be cleaned every _____ hours.

| a. | 24 | c. | 72 |
| --- | --- | --- | --- |
| b. | 48 | d. | 96 |

ANS:	A

The infant flow sensor should be cleaned every 24 hours to maintain accurate tidal volume measurements and flow-triggering capabilities.

PTS:	1	REF:	Page 480, Box 14-8

37.	Which of the following statements concerning the CareFusion Bird Partner IIi monitor is true?

1. This device uses a vortex ultrasonic device to measure flow.

2. To prevent obstructions, 12 mL/min of gas is passed through the pressure line.

3. This device can measure artificial airway leak only with the use of the infant sensor.

4. Placement of the pediatric and infant sensors is the same in the ventilator tubing.

| a. | 1 and 2 | c. | 3 and 4 |
| --- | --- | --- | --- |
| b. | 2 and 3 | d. | 1 and 4 |

ANS:	B

The gas inlet on the back of the ventilator connects to a 50-psig source and is used to inject 12 mL/min of gas through the pressure line to prevent obstructions within the line and to prevent water from entering the differential pressure transducer. By pressing and holding the “CONT V” button on the touch pad, the clinician enables real-time flow through the infant sensor to be displayed. The clinician can use the baseline leak or the flow displayed between breaths to determine the best trigger sensitivity setting and to prevent auto-triggering. This device uses a variable-orifice differential pressure transducer.

PTS:	1	REF:	Page 480

38.	When using the CareFusion V.I.P. Bird Gold Infant/Pediatric Ventilator with the infant flow sensor, the flow display comp button is pressed. The ventilator’s monitor display window shows real-time inspiratory and expiratory flows of 9 L/min and 8.2 L/min. The most appropriate assist sensitivity control setting is ____ L/min.

| a. | 0.3 | c. | 0.8 |
| --- | --- | --- | --- |
| b. | 0.5 | d. | 1.0 |

ANS:	D

The baseline leak is 9 – 8.2 = 0.8 L/min. The assist sensitivity control should be set at least 0.2 L/min above the detected leak, which is 0.8 L/min + 0.2 L/min = 1.0 L/min.

PTS:	1	REF:	Page 480

39.	A pediatric patient is being ventilated with a CareFusion V.I.P. Bird Gold Infant/Pediatric Ventilator in the assist/control volume mode with appropriately set trigger sensitivity. Physical assessment of the patient reveals the use of accessory muscles during patient triggering. What is the most appropriate action?

| a. | Turn the sensitivity off. |
| --- | --- |
| b. | Switch to another ventilator. |
| c. | Turn the leak compensation off. |
| d. | Remove the source of the airway leak. |

ANS:	C

Leak compensation is used when an artificial airway leak prevents the ventilator from otherwise maintaining the positive end-expiratory pressure level. Small increments of flow are introduced into the circuit to provide back-pressure compensation for the leak. When leak compensation is in use, the patient’s ability to trigger may be diminished.

PTS:	1	REF:	Page 486

40.	The bias flow for the CareFusion V.I.P. Gold Ventilator when the Infant “Smart” Flow Sensor is in use is ____ L/min.

| a. | 2 | c. | 4 |
| --- | --- | --- | --- |
| b. | 3 | d. | 5 |

ANS:	B

When the infant sensor is in use, the bias flow is automatically turned on at 3 L/min.

PTS:	1	REF:	Page 485

41.	The bias flow for the CareFusion V.I.P. Gold Ventilator when the Pediatric “Smart” Flow Sensor is in use is _____ L/min.

| a. | 3 | c. | 5 |
| --- | --- | --- | --- |
| b. | 4 | d. | 6 |

ANS:	C

When the Pediatric “Smart” Flow Sensor is in use, bias flow operates at a fixed 5 L/min flow when it is turned on.

PTS:	1	REF:	Page 481

42.	Which of the following ventilatory modes is available only on the CareFusion V.I.P. Bird Gold Infant Ventilator and not on the other V.I.P. Infant ventilators?

| a. | Time-cycled, pressure-limited |
| --- | --- |
| b. | Volume-assured pressure support |
| c. | Assist/control volume-cycled |
| d. | Synchronized intermittent mandatory ventilation/continuous positive airway pressure volume-cycled |

ANS:	B

The volume-assured pressure support mode is available only on the CareFusion V.I.P. Bird Gold model.

PTS:	1	REF:	Page 481

43.	When volume-assured pressure support is selected in the    synchronized intermittent mandatory ventilation/continuous positive airway pressure/pressure support setting on the CareFusion    V.I.P. Bird Gold Infant Ventilator, mandatory breaths are classified as which of the following?

| a. | Pressure-controlled, time-cycled |
| --- | --- |
| b. | Volume-controlled, flow-cycled |
| c. | Pressure-controlled, volume-guaranteed |
| d. | Volume-controlled, pressure-guaranteed |

ANS:	C

When volume-assured pressure support is selected from the synchronized intermittent mandatory ventilation/continuous positive airway pressure/pressure support column, mandatory pressure-controlled breaths are delivered at a guaranteed volume.

PTS:	1	REF:	Page 483

44.	On the CareFusion V.I.P. Bird Gold Infant Ventilator, the control that ends a pressure support breath in case of a leak also does which of the following to the breath?

| a. | Flow-cycle | c. | Volume-cycle |
| --- | --- | --- | --- |
| b. | Time-cycle | d. | Pressure-cycle |

ANS:	B

The pressure support/volume-assured pressure support time limit control will time-cycle the normally flow-cycled pressure support breath when there is a leak.

PTS:	1	REF:	Page 483

45.	Which of the following CareFusion V.I.P. Bird Gold Infant Ventilator controls can adjust pressure delivery at the beginning of inspiration in the pressure control, volume-assured pressure support, or pressure support mode?

| a. | Bias flow | c. | Rise time |
| --- | --- | --- | --- |
| b. | Assist sensitivity | d. | Flow display comp |

ANS:	C

The rise time control adjusts the inspiratory pressure rise time. A setting of 1 is the fastest rise time, and a setting of 7 is the slowest rise time. The bias flow control is always present when the Infant “Smart” Sensor is used and can be turned on or off when the Pediatric “Smart” Sensor is used. The assist sensitivity controls how sensitive the ventilator is to the patient’s efforts, and the flow display comp is used to display real-time inspiratory and expiratory flows.

PTS:	1	REF:	Pages 483-484

46.	When the Dräger Babylog 8000+ is used in the pressure support mode, the patient’s peak inspiratory flow is 30 L/min. At what flow rate will the ventilator cycle to expiration?

| a. | 3.0 L/min | c. | 6.0 L/min |
| --- | --- | --- | --- |
| b. | 4.5 L/min | d. | 7.0 L/min |

ANS:	B

Pressure support breaths are flow-cycled on the Dräger Babylog 8000+ at 15% of peak flow.

PTS:	1	REF:	Page 486

47.	On a Dräger Babylog 8000, which of the following variables can trigger mandatory breaths in the synchronized intermittent mandatory ventilation mode?

| a. | Volume and flow | c. | Time and volume |
| --- | --- | --- | --- |
| b. | Pressure and flow | d. | Pressure and time |

ANS:	C

The mandatory breaths in the synchronized intermittent mandatory ventilation mode on the Dräger Babylog 8000 are either time-triggered or patient-triggered. Patient-triggering is accomplished by volume, when a patient’s spontaneous inspiratory volume is equal to or greater than the set trigger volume.

PTS:	1	REF:	Page 487

48.	The oxygen concentration alarm limits on the Dräger Babylog 8000 are set _____ at _____ of the setting.

| a. | by the operator; ± 5% | c. | automatically; ± 5% |
| --- | --- | --- | --- |
| b. | by the operator; ± 4% | d. | automatically; ± 4% |

ANS:	D

The Dräger Babylog 8000 microprocessor automatically sets the oxygen concentration limits at ± 4%.

PTS:	1	REF:	Page 488

49.	On the Dräger Babylog 8000, which of the following is an operator-adjustable alarm?

| a. | Apnea |
| --- | --- |
| b. | Oxygen concentration |
| c. | Loss of positive end-expiratory pressure/continuous positive airway pressure |
| d. | High-pressure limit |

ANS:	A

Adjustable alarm limits include apnea (5 to 20 seconds), high and low minute ventilation, and minute ventilation delay.

PTS:	1	REF:	Page 488

50.	Which operating mode on the Dräger Babylog 8000 allows the clinician to adjust the flow during the expiratory phase to match the patient’s needs during mandatory and spontaneous breaths?

| a. | Assist/control |
| --- | --- |
| b. | Intermittent mandatory ventilation |
| c. | Variable inspiratory, variable expiratory |
| d. | Continuous positive airway pressure |

ANS:	C

The variable inspiratory, variable expiratory operating mode allows the clinician to adjust the flow during the expiratory phase to match the patient’s needs during mandatory and spontaneous breaths.

PTS:	1	REF:	Page 488

51.	Which of the following statements concerning the volume guarantee mode on the Dräger Babylog 8000 ventilator is true?

| a. | The breaths are volume-limited in this mode. |
| --- | --- |
| b. | This mode does not operate with assist/control. |
| c. | The Dräger Babylog 8000 and 8000+ both have this mode. |
| d. | Positional endotracheal tube leaks may be hazardous with this mode. |

ANS:	D

When the volume guarantee mode is activated, the ventilator continuously measures and compares inspired and expired tidal volume. Each following breath uses the comparisons from the previous breath to make adjustments to the peak inspiratory pressure so that a volume as close as possible to the preset value can be delivered. If there is a positional leak and it is active, the expired tidal volume will be lower than the inspired tidal volume, and the ventilator will increase the pressure to attempt to match the two monitored volumes. This increase in pressure continues if the leak continues. However, if the patient’s position changes and the leak is not occurring, the breath delivered will be significantly larger than intended. The limit to the amount of volume is 130% of the targeted volume. At this point, the exhalation valve will open, and no additional ventilator-driven gas will be delivered to the patient. Close monitoring of patient-ventilator interaction is needed during use of this mode.

PTS:	1	REF:	Page 490

52.	A trigger sensitivity setting of 1 on a Dräger Babylog 8000 Infant Ventilator indicates ____ trigger.

| a. | 1 L/min flow | c. | 0.25 L/min flow |
| --- | --- | --- | --- |
| b. | 1 mL volume | d. | 1.5 mL volume |

ANS:	C

The trigger sensitivity range is 1 to 10, with 1 (minimum) representing increased trigger sensitivity and 10 (maximum) representing the least-sensitive trigger. A setting of 1 indicates that when the Dräger Babylog measures a flow change of 0.25 L/min (straight flow-trigger), the mandatory breath is synchronized with the patient’s effort.

PTS:	1	REF:	Page 488

53.	Mean airway pressure can be increased when ventilating with the Bunnell Life Pulse High-Frequency Jet Ventilator by doing which of the following?

| a. | Closing off the purge valve |
| --- | --- |
| b. | Decreasing the rate on the Bunnell |
| c. | Decreasing the positive end-expiratory pressure on the Bunnell |
| d. | Increasing rate and volume of sigh breaths on the conventional ventilator |

ANS:	D

The function of the purge valve is to provide a moisture-free environment in the monitoring line of the hi-lo jet tube. Decreasing the rate decreases the mean airway pressure because less time is spent at the inspiratory pressure. Decreasing the positive end-expiratory pressure level decreases the mean airway pressure. Increasing the rate and volume of the sigh breaths increases the mean airway pressure.

PTS:	1	REF:	Page 494

54.	The “cannot meet PIP” alarm is active on a Bunnell Life Pulse High-Frequency Jet Ventilator. The respiratory therapist should do which of the following?

1. Check the humidifier cartridge for a leak.

2. Check the jet tube connection to the circuit.

3. Look for kinks in the jet tube connections.

4. Check the dump valve for proper function.

| a. | 2 | c. | 3 and 4 |
| --- | --- | --- | --- |
| b. | 1 and 2 | d. | 1, 2, and 3 |

ANS:	D

The “cannot meet PIP” alarm can be a result of a leak in the humidifier cartridge/patient circuit; an incomplete connection of the circuit to the jet tube; a defective or damaged jet tube (e.g., a kinked tube, improper positioning, occlusion, leak); the present settings being insufficient to ventilate larger patients; a patient fighting the Life Pulse; or an ineffective pinch valve opening action resulting in higher servo pressures being necessary to meet current settings.

PTS:	1	REF:	Page 493

55.	The Bunnell Life Pulse High Frequency Jet Ventilator creates high-frequency jet ventilation by using which of the following mechanisms?

| a. | Rotary piston | c. | Spinning ball |
| --- | --- | --- | --- |
| b. | Pinch valve | d. | Flow interrupter |

ANS:	B

An electromagnetic solenoid activates the pinch valve, which breaks the flow of pressurized gas into small bursts with the pinch-and-release action of the silicone tube of the patient breathing circuit.

PTS:	1	REF:	Page 494

56.	The Bunnell Life Pulse High-Frequency Jet Ventilator is being set up to ventilate an infant who has a regular endotracheal tube with a 3.5-mm internal diameter Life Port adapter. Conventional peak inspiratory pressure is set at 30 cm H2O. The high-frequency jet ventilation peak inspiratory pressure should be set at a pressure of _____ cm H2O.

| a. | 25.5 | c. | 28.5 |
| --- | --- | --- | --- |
| b. | 27.0 | d. | 30.0 |

ANS:	B

When a LifePort adapter larger than 2.5-mm internal diameter is used, the clinician should set up the initial high-frequency jet ventilation peak inspiratory pressure at 90% of the conventional peak inspiratory pressure. Therefore, 90% of 30 cm H2O is 27 cm H2O.

PTS:	1	REF:	Page 495

57.	The Bunnell Life Pulse High-Frequency Jet Ventilator is being set up to ventilate an infant who has already been intubated with a regular endotracheal tube. A 2.5-mm internal diameter Life Port adapter is being used. The conventional peak inspiratory pressure is set at 26 cm H2O. The high-frequency jet ventilation peak inspiratory pressure should be set at a pressure of _____ cm H2O.

| a. | 23.4 | c. | 26 |
| --- | --- | --- | --- |
| b. | 24.7 | d. | 28 |

ANS:	C

When using a Life Port adapter with a 2.5-mm internal diameter, the clinician should set up the initial high-frequency jet ventilation peak inspiratory pressure equal to the conventional peak inspiratory pressure. Therefore, 26 cm H2O should be the setting.

PTS:	1	REF:	Page 495

58.	What is the function of the structure labeled “A” in the following figure of a Bunnell Life Pulse High-Frequency Jet Ventilator?

| a. | It measures volume. | c. | It delivers the jet bursts. |
| --- | --- | --- | --- |
| b. | It monitors pressure. | d. | It allows conventional ventilation. |

ANS:	B

“A” is the monitoring lumen of the hi-lo tube and is used to monitor pressures at the distal end of the hi-lo jet tube.

PTS:	1	REF:	Page 495, Figure 14-32

59.	The respiratory therapist has determined that a 4.0-mm internal diameter standard endotracheal tube is necessary for the patient. What is the equivalent hi-lo jet tube to be used with a Bunnell Life Pulse High-Frequency Jet Ventilator?

| a. | 3.0-mm internal diameter hi-lo jet tube |
| --- | --- |
| b. | 3.5-mm internal diameter hi-lo jet tube |
| c. | 4.0-mm internal diameter hi-lo jet tube |
| d. | 4.5-mm internal diameter hi-lo jet tube |

ANS:	B

The external diameter of a hi-lo jet tube is approximately equal to the external diameter of a standard endotracheal tube that is a half-size larger. Therefore, hi-lo tube size = standard endotracheal tube size – 0.5 mm; 4.0 mm is standard endotracheal tube size, so 4.0 mm – 0.5 = 3.5-mm hi-lo tube size.

PTS:	1	REF:	Page 495

60.	At what point in the Bunnell Life Pulse High-Frequency Jet Ventilator, shown in the following figure, should a bacterial filter be placed?

| a. | A | c. | C |
| --- | --- | --- | --- |
| b. | B | d. | D |

ANS:	D

The bacterial filter should be placed downstream from the pinch valve to provide particle filtration.

PTS:	1	REF:	Page 494, Figure 14-30

61.	What is the function of the purge valve on the Bunnell Life Pulse High-Frequency Jet Ventilator?

| a. | It releases internal pressure. |
| --- | --- |
| b. | It allows water into the humidifier cartridge. |
| c. | It opens to dry out the monitoring line of the hi-lo jet tube. |
| d. | It breaks pressurized gas flow into small bursts of gas. |

ANS:	C

The purge port supplies gas to the purge valve, which is located in the patient box. The gas is used to provide a moisture-free environment in the monitoring line of the hi-lo jet tube.

PTS:	1	REF:	Page 494

62.	The jet rate range for the Bunnell Life Pulse High-Frequency Jet Ventilator is _____ insufflations/min.

| a. | 60 to 150 | c. | 240 to 660 |
| --- | --- | --- | --- |
| b. | 150 to 240 | d. | 660 to 900 |

ANS:	C

The Life Pulse High-Frequency Jet Ventilator can provide 240 to 660 insufflations/min. The range 660 to 900 insufflations/min is considered high-frequency oscillation. The range 60 to 150 insufflations/min is considered high-frequency positive pressure ventilation.

PTS:	1	REF:	Page 493

63.	The “jet valve fault” alarm is active on a Bunnell Life Pulse High-Frequency Jet Ventilator. This problem can be corrected by doing which of the following?

| a. | Changing out the pinch valve tubing |
| --- | --- |
| b. | Changing the tubing to the dump valve outlet |
| c. | Reconnecting the purge valve to the patient box |
| d. | Reconnecting the pressure-monitoring tube to the patient box |

ANS:	A

The jet valve fault alarm alerts the clinician that the pinch valve in the patient box is not functioning appropriately. This could be because of a worn-out pinch valve tube.

PTS:	1	REF:	Page 493

64.	How many cycles per minute is 15 Hz?

| a. | 300 | c. | 900 |
| --- | --- | --- | --- |
| b. | 600 | d. | 1200 |

ANS:	C

A value of 1 Hz is 60 cycles/min. Therefore, 15 Hz  60 cycles/min = 900 cycles/min.

PTS:	1	REF:	Page 498

65.	Which of the following actions will maximize the life of the oscillator mechanism on a CareFusion 3100A High-frequency Oscillatory Ventilator?

| a. | Use low-bias flow settings. |
| --- | --- |
| b. | Minimize time and amplitude settings. |
| c. | Maintain the piston at the center of the bar graph. |
| d. | Keep the piston toward the maximum inspiratory limit. |

ANS:	C

Keeping the piston in the center of the bar graph will prevent the piston from being driven against a mechanical stop for an extended period. Operating against a mechanical stop will wear out the oscillator mechanism quickly.

PTS:	1	REF:	Page 500

66.	Increasing the bias flow on the CareFusion 3100A High-frequency Oscillatory Ventilator would cause which of the following to occur?

| a. | Decrease in amplitude | c. | Increase in mean airway pressure |
| --- | --- | --- | --- |
| b. | Increase in percent inspiratory time | d. | Automatic centering of piston |

ANS:	C

Changes to the bias flow can increase mean airway pressure. A limit to how much mean airway pressure can develop in the system is controlled by the mean pressure limit control knob. This control is used to protect the patient from inadvertent increases in mean airway pressure caused by changes in bias flow, percent inspiratory time, or frequency.

PTS:	1	REF:	Page 498

67.	After an increase in the bias flow on a CareFusion 3100A, the airway pressure has increased. What is the most appropriate action?

| a. | Decrease the airway pressure. | c. | Re-center the piston. |
| --- | --- | --- | --- |
| b. | Decrease the power. | d. | Maintain the current airway pressure. |

ANS:	A

If a change in the airway pressure occurs because another control has been adjusted, the mean pressure adjust control knob should be used to return the airway pressure to the desired level.

PTS:	1	REF:	Page 498

68.	Which of the following statements concerning the CareFusion 3100A and 3100B is true?

| a. | There is a mean pressure limit control on both models. |
| --- | --- |
| b. | Piston centering is automatically performed by electronic sensors. |
| c. | The high mean airway pressure alarm is the same on both models. |
| d. | The inspiratory time control limits have been expanded on the 3100B. |

ANS:	B

The CareFusion 3100B has automatic piston-centering, and the center piston control has been eliminated from the 3100B. The 3100A has a mean pressure limit control; on the 3100B, this control has been eliminated. The high mean airway pressure alarm is 50 cm H2O on the 3100A and 60 cm H2O on the 3100B. The inspiratory time control limits are the same on both models.

PTS:	1	REF:	Page 500

69.	Lowering the frequency on the CareFusion 3100A or 3100B would do which of the following?

1. Decrease the volume delivered

2. Increase the inspiratory time

3. Increase the volume delivered

4. Increase the mean airway pressure

| a. | 1 and 2 | c. | 3 and 4 |
| --- | --- | --- | --- |
| b. | 2 and 3 | d. | 1 and 4 |

ANS:	C

When the frequency is lowered, the piston is allowed more travel time. This results in greater piston displacement, or more delivered volume, the amount of which is unknown.

PTS:	1	REF:	Page 499, Box 14-10

70.	An increase in % inspiratory time on the CareFusion 3100A would cause which of the following?

1. Alteration in the symmetry of the oscillator waveform

2. Decrease in the frequency

3. Increase in the airway pressure

4. Decrease in volume

| a. | 1 and 2 | c. | 1 and 3 |
| --- | --- | --- | --- |
| b. | 2 and 4 | d. | 3 and 4 |

ANS:	C

An increase in % inspiratory time would cause more time to be spent in the inspiratory phase, so the symmetry of the waveform would be altered. The change to % inspiratory time would also cause an increase in airway pressure. This change should be readjusted by using the power control.

PTS:	1	REF:	Page 498

71.	The purpose of the Venturi-type air amplifier on the CareFusion 3100A is to do which of the following?

| a. | Provide gas for bias flow |
| --- | --- |
| b. | Add gas to increase the volume delivered |
| c. | Cool the electrical coil that moves the piston |
| d. | Prevent the square wave driver from overheating |

ANS:	C

The function of the Venturi-type air amplifier is to introduce cooling air around the electrical coil.

PTS:	1	REF:	Page 496

72.	When helium/oxygen (heliox) is used with the CareFusion AVEA Ventilator in the infant/pediatric modes, which of the following statements is true?

| a. | The hot-wire flow sensor must be used. |
| --- | --- |
| b. | Heliox may not be used with this ventilator. |
| c. | The ventilator’s internal flow sensors can be used. |
| d. | Data and graphics must be adjusted by the operator. |

ANS:	C

With heliox use, the CareFusion AVEA’s internal inspiratory flow sensor or the VarFlex flow sensor may be used.

PTS:	1	REF:	Page 502

73.	During infant ventilation with the CareFusion AVEA Ventilator and the Cosmo Plus Monitor, the readings on the Cosmo Plus Monitor have “blanked out.” What is the cause of this problem?

| a. | Elevated pressure in the device | c. | Monitor is out of calibration |
| --- | --- | --- | --- |
| b. | Condensation in the device | d. | Monitor is calibrating |

ANS:	B

One of the concerns with the use of the Cosmo Plus is that the readings become “blanked out” when water condensation accumulates in the device.

PTS:	1	REF:	Page 502

74.	A patient who weighs 4000 g is being ventilated by using a CareFusion AVEA Ventilator. How much volume will be lost in the tubing when the pressure is set at 15 cm H2O?

| a. | 8 mL | c. | 15 mL |
| --- | --- | --- | --- |
| b. | 10 mL | d. | 18 mL |

ANS:	C

When using the CareFusion AVEA with a patient weighing 5 kg or less, an infant circuit with a compressible volume of approximately 1 mL/cm H2O should be used. This will cause 15 mL to be left in the tubing during inspiration.

PTS:	1	REF:	Page 502

75.	NeoFlow on a Dräger E-4 requires the installation of a flow sensor into which part of the patient circuit?

| a. | Before the inspiratory line |
| --- | --- |
| b. | Between the endotracheal tube and the Y-connector |
| c. | Proximal to the exhalation flow sensor |
| d. | Between the inspiratory filter and the humidifier |

ANS:	B

NeoFlow requires the placement of a flow sensor between the artificial airway and the Y-connector of the patient circuit.

PTS:	1	REF:	Page 504

76.	When NeoFlow on the Dräger E-4 is activated, how much bias flow is initiated?

| a. | 2 L/min | c. | 6 L/min |
| --- | --- | --- | --- |
| b. | 4 L/min | d. | 8 L/min |

ANS:	C

The base, or bias, flow during NeoFlow use is 6 L/min.

PTS:	1	REF:	Page 504

77.	NeoFlow on the Dräger E-4 is essential for infants with which of the following weights?

1. 3.8 kg

2. 5.0 kg

3. 7.6 kg

4. 8.4 kg

| a. | 1 only | c. | 2, 3, and 4 |
| --- | --- | --- | --- |
| b. | 1 and 2 | d. | 1, 2, 3, and 4 |

ANS:	B

NeoFlow is essential with patients weighing less than 6 kg.

PTS:	1	REF:	Page 504

78.	On a Dräger E-4 operating in NeoFlow, which of the following trigger ranges is recommended for neonatal and pediatric ventilation?

| a. | 0.1 to 0.5 L/min | c. | 0.5 to 3.5 L/min |
| --- | --- | --- | --- |
| b. | 0.3 to 3 L/min | d. | 1 to 4 L/min |

ANS:	B

A trigger range of 0.3 to 3 L/min is recommended for neonatal and pediatric ventilation.

PTS:	1	REF:	Page 504

79.	While a patient is being ventilated with NeoFlow on a Dräger E-4, an audible and visual alarm is active, and all breaths are being time-triggered. What is the most appropriate action for the respiratory therapist to take?

| a. | Suction the endotracheal tube | c. | Dry out the flow sensor |
| --- | --- | --- | --- |
| b. | Change the ventilator out | d. | Switch to AutoFlow |

ANS:	C

If the NeoFlow sensor becomes partially occluded with water (or secretions), it will cease operation and trigger audible and visual alarms. When this occurs, the ventilator will continue to operate in the pressure-control mode.

PTS:	1	REF:	Page 504

80.	Which of the following statements concerning the use of AutoFlow for neonates on the Dräger E-4 is true?

| a. | AutoFlow is an adult mode and should not be used with neonates. |
| --- | --- |
| b. | The ventilator’s internal flow sensor is always used with neonates on AutoFlow. |
| c. | Two test breaths determine the appropriate pressure to deliver volume in AutoFlow. |
| d. | In AutoFlow, the highest pressure to deliver the volume is the high-pressure limit setting. |

ANS:	C

When the flow sensor is attached during AutoFlow in neonates, the unit provides two test breaths, first with a pressure of 5 cm H2O and then with a pressure equal to 75% of that required to deliver the set volume.

PTS:	1	REF:	Page 505

81.	Use of the Dräger E-4 nebulizer in the “Neo” setting will do which of the following?

| a. | Increase the base flow from 6 L/min to 11 L/min |
| --- | --- |
| b. | Cause a 6% decrease in the oxygen concentration in any mode |
| c. | Divert 6 L/min away from the patient circuit to the nebulizer |
| d. | Decrease oxygen concentration significantly with rates greater than 12 breaths/min |

ANS:	D

Use of the Dräger E-4 nebulizer in the “Neo” setting increases the base flow from 6 L/min to 9 L/min. Delivered oxygen concentrations decrease slightly from the level set. The amount of decrease is less than 4% at breath rates greater than 12 per minute. When operating at mandatory breath rates of less than 12 per minute, the nebulizer should not be used because of the potentially significant drop in oxygen concentration.

PTS:	1	REF:	Page 505

82.	In the neonatal option, the Maquet Servoi Ventilator has a bias flow of _____ L/min.

| a. | 0.5 | c. | 1.5 |
| --- | --- | --- | --- |
| b. | 1.0 | d. | 2.0 |

ANS:	A

The bias flow is automatically 0.5 L/min with the neonatal option.

PTS:	1	REF:	Page 505

83.	In the neonatal option, the Maquet Servoi Ventilator has a maximum available flow of _____ L/min.

| a. | 28 | c. | 33 |
| --- | --- | --- | --- |
| b. | 30 | d. | 35 |

ANS:	C

The maximum flow available with the neonatal option is 33 L/min.

PTS:	1	REF:	Page 505

84.	The range of noninvasive ventilation nasal continuous positive airway pressure on the Maquet Servoi Ventilator is _____ cm H2O.

| a. | 1 to 10 | c. | 3 to 25 |
| --- | --- | --- | --- |
| b. | 2 to 20 | d. | 4 to 30 |

ANS:	B

During nasal continuous positive airway pressure, the ventilator provides enough flow to maintain the pressure settings of 2 to 20 cm H2O.

PTS:	1	REF:	Page 507

85.	There is a 3-L/min leak around the nasal prongs during nasal continuous positive airway pressure on the Maquet Servoi Ventilator. This will cause which of the following?

| a. | Bias flow will drop to 7.5 L/min. |
| --- | --- |
| b. | Disconnect alarm will become active. |
| c. | Flow will automatically increase. |
| d. | Expiratory sensitivity will decrease. |

ANS:	C

If a leak occurs around the prongs, the ventilator will automatically and immediately increase the flow available to maintain the pressure.

PTS:	1	REF:	Page 505

86.	What is the minimum ideal body weight for ventilating a pediatric patient with a Puritan Bennett 840 without the NeoMode option?

| a. | 4 kg | c. | 6 kg |
| --- | --- | --- | --- |
| b. | 5 kg | d. | 7 kg |

ANS:	D

Without the NeoMode option, this ventilator is designed for use with pediatric patients who have an ideal body weight of at least 7 kg.

PTS:	1	REF:	Page 508

87.	What is the minimum ideal body weight for ventilating an infant with a Puritan Bennett 840 with the NeoMode option?

| a. | 0.5 kg | c. | 1.5 kg |
| --- | --- | --- | --- |
| b. | 1 kg | d. | 2 kg |

ANS:	A

Infants weighing 0.5 to 7 kg can be ventilated by using the NeoMode option on the Puritan Bennett 840.

PTS:	1	REF:	Page 508

88.	Heliox delivery is available through which of the following ventilators?

| a. | Maquet Servoi | c. | Dräger Babylog 8000+ |
| --- | --- | --- | --- |
| b. | CareFusion V.I.P. Bird | d. | CareFusion AVEA |

ANS:	D

Heliox delivery is available in the CareFusion AVEA Comprehensive model or can be added as an option in the other three    models.

PTS:	1	REF:	Pages 501-502

89.	The pressure high flowmeter is set at 12 L/min; which of the following flow rates will be delivered?

| a. | 10.2 to 13.8 L/min | c. | 11 to 14.5 L/min |
| --- | --- | --- | --- |
| b. | 10.5 to 14.2 L/min | d. | 11 to 15 L/min |

ANS:	A

The flowmeter is adjustable from 0 to 15 L/min with an accuracy of ± 15% of the selected value.

PTS:	1	REF:	Page 466, Table 14-1

90.	A CareFusion SiPAP device is being used to support an infant; the pressure is set a high pressure of 12 cm H2O and a low pressure of 4 cm H2O. The low airway pressure alarm will be activated if the pressure drops _______ cm H2O below the set level.

| a. | 2 | c. | 6 |
| --- | --- | --- | --- |
| b. | 4 | d. | 8 |

ANS:	A

The low airway pressure alarm is activated if the pressure falls 2 cm H2O below the set pressure.

PTS:	1	REF:	Page 468, Table 14-2

91.	What is the available range for peak inspiratory pressure for a Dräger Babylog 8000 infant ventilator?

| a. | 5 to 50 cm H2O | c. | 10 to 80 cm H2O |
| --- | --- | --- | --- |
| b. | 10 to 60 cm H2O | d. | 10 to 100 cm H2O |

ANS:	C

The available range for the peak inspiratory pressure for the Dräger Babylog 8000 Infant Ventilator is 10 to 80 cm H2O.

PTS:	1	REF:	Page 488, Table 14-6

92.	What is the available range for setting the tidal volume on the CareFusion AVEA Ventilator?

| a. | 1 to 150 mL | c. | 1 to 250 mL |
| --- | --- | --- | --- |
| b. | 1 to 200 mL | d. | 1 to 300 mL |

ANS:	D

See Table 14-7 for answer.

Page 14-124, Table 14-7

PTS:	1	REF:	Page 503, Table 14-7

93.	What range can be used for adjustment of peak flow when the CareFusion Avea is used for pediatric patients?

| a. | 1 to 50 L/min | c. | 1 to 100 L/min |
| --- | --- | --- | --- |
| b. | 1 to 75 L/min | d. | 1 to 125 L/min |

ANS:	B

The peak flow for the Avea can be adjusted from 1 to 75 L/min when the ventilator is used in the pediatric mode.

PTS:	1	REF:	Page 503, Table 14-7

94.	What is the available range for setting the respiratory rate on the Covidien PB 840 Ventilator when it is used in the NeoMode?

| a. | 1-100 breaths/min | c. | 1-150 breaths/min |
| --- | --- | --- | --- |
| b. | 1-125 breaths/min | d. | 1-175 breaths/min |

ANS:	C

The Covidien PB 840 Ventilator can deliver a respiratory rate of 1 to 150 breaths/min.

PTS:	1	REF:	Page 508, Table 14-8

95.	Which of the following events occur when the mean airway pressure exceeds 50 cm H2O?

| a. | The oscillator decreases by 50%. |
| --- | --- |
| b. | The pinch valve closes to decrease pressure. |
| c. | The dump valve opens, causing the oscillator to stop. |
| d. | The mean airway pressure is reduced by 25%. |

ANS:	C

The mean airway pressure alarm greater than 50 cm H2O is activated if the mean airway pressure rises above 50 cm H2O for any reason. When this alarm is activated, the dump valve opens, causing the oscillator to stop.

PTS:	1	REF:	Page 499, Box 14-10

96.	When using the Dräger Evita XL, which of the following modes allow for the use of AutoFlow when the NeoFlow sensor is being used?

1. Continuous mandatory ventilation

2. Volume-assured pressure support

3. Synchronized intermittent mandatory ventilation

4. Mandatory minute ventilation

| a. | 1 only | c. | 1, 3, and 4 |
| --- | --- | --- | --- |
| b. | 1 and 2 | d. | 1, 2, 3, and 4 |

ANS:	C

With the sensor in place, AutoFlow functions in the continuous mandatory ventilation, synchronized intermittent mandatory ventilation, and mandatory minute ventilation modes.

PTS:	1	REF:	Page 504