Chapter 04: Administering Medical Gases: Regulators, Flowmeters, and

MULTIPLE CHOICE

1.	What device helps reduce high-pressure gases from cylinders or bulk storage units to a lower working pressure?

| a. | Regulator | c. | Bourdon gauge |
| --- | --- | --- | --- |
| b. | Kinetic tube | d. | Thorpe tube |

ANS:	A

The regulator is the only device in which pressure is regulated. The other devices relate to flow.

PTS:	1	REF:	Page 84

2.	The chamber located above the diaphragm of a single-stage regulator is which of the following?

| a. | Lower chamber | c. | High-pressure chamber |
| --- | --- | --- | --- |
| b. | Pressure-relief valve | d. | Adjustable regulator |

ANS:	C

The high-pressure chamber is located above the diaphragm and contains a valve stem that will close the inlet when the pressure reaches its setting.

PTS:	1	REF:	Page 84

3.	The gas flow into the high-pressure side regulator is dependent on what two opposing forces?

1. Gas pressure above the diaphragm

2. Gas pressure at the diaphragm

3. Spring tension below the diaphragm

4. Spring tension above the diaphragm

| a. | 1 and 3 | c. | 2 and 4 |
| --- | --- | --- | --- |
| b. | 2 and 3 | d. | 1 and 4 |

ANS:	A

When the force offered by the high-pressure gas above the diaphragm equals the force offered by spring tension, the diaphragm is straight and the inlet valve is closed.

PTS:	1	REF:	Page 84

4.	Delivered pressures on common adjustable regulators are between _____ and ____ pounds per square inch gauge (psig).

| a. | 0; 50 | c. | 50; 100 |
| --- | --- | --- | --- |
| b. | 0; 100 | d. | 50; 200 |

ANS:	B

Most adjustable regulators can be set to deliver pressure between 0 and 100 psig.

PTS:	1	REF:	Page 84

5.	Which regulator supports the theory that gas pressure is gradually reduced as gas flows from a high-pressure source through a series of stages to the outlet?

| a. | Single-stage regulator | c. | Preset regulator |
| --- | --- | --- | --- |
| b. | Flowmeter | d. | Multistage regulator |

ANS:	D

Gas from a compressed cylinder (e.g., 2200 psig) enters the first stage of a two-stage regulator, and the gas pressure is reduced to an intermediate pressure (e.g., 700 psig). This lower pressure gas then enters into the second stage of the regulator, where the gas pressure is further reduced to the desired working pressure (e.g., 50 psig) before the gas reaches the outlet.

PTS:	1	REF:	Page 85

6.	Which device is more precise in terms of controlling gas pressure?

| a. | Bourdon flowmeter | c. | Multistage regulator |
| --- | --- | --- | --- |
| b. | Thorpe tube flowmeter | d. | Flow restrictors |

ANS:	C

Multistage regulators can control gas pressure with more precision than single-stage regulators because the pressure is gradually reduced.

PTS:	1	REF:	Page 85

7.	The most commonly used flowmeter in respiratory care is which of the following?

| a. | Thorpe tube flowmeter | c. | Flow restrictor |
| --- | --- | --- | --- |
| b. | Bourdon gauge | d. | Reducing valve |

ANS:	A

Thorpe tubes are the most commonly used flowmeters in respiratory care.

PTS:	1	REF:	Page 86

8.	The problem most often encountered with a Thorpe tube flowmeter is:

| a. | patient outlet becomes occluded. |
| --- | --- |
| b. | gas leaks caused by faulty valve seats. |
| c. | a cracked tube. |
| d. | prolonged heat exposure. |

ANS:	B

This problem is usually detected when the flowmeter is turned off completely, but gas can be heard continuing to flow from the flowmeter outlet; in this case, the flowmeter should be replaced. This problem can also occur because of constant wear and tear.

PTS:	1	REF:	Page 87

9.	Pressure-compensated flowmeters provide which of the following?

| a. | Inaccurate estimates of flow | c. | Back pressure |
| --- | --- | --- | --- |
| b. | Accurate estimates of flow | d. | A means to estimate leaks |

ANS:	B

On pressure-compensated flowmeters, the needle valve controlling gas flow out of the flowmeter is located distal to the Thorpe tube. This arrangement allows the pressure in the indicator tube to be maintained at the source gas pressure, which ensures an accurate estimate of flow.

PTS:	1	REF:	Page 86

10.	A float that jumps and falls to zero when the needle valve is closed and is attached to a bulk system is indicative of a _____ flowmeter.

| a. | crack in the | c. | non-compensated |
| --- | --- | --- | --- |
| b. | Bourdon | d. | pressure-compensated |

ANS:	D

This floating movement occurs because the source gas must pass through the indicator tube before it reaches the needle valve.

PTS:	1	REF:	Page 86

11.	Which device uses the following principle? As the driving pressure is increased, the flow from the flowmeter outlet increases.

| a. | Compensated Thorpe tube flowmeter |
| --- | --- |
| b. | Fixed-orifice flowmeter |
| c. | Bourdon flowmeter |
| d. | Non-compensated pressure flowmeter |

ANS:	C

A Bourdon flowmeter is actually a reducing valve that controls the pressure gradient across an outlet with a fixed orifice.

PTS:	1	REF:	Page 87

12.	What is (are) the type(s) of flow restrictor(s)?

1. Fixed-orifice

2. Adjustable, multiple-orifices

3. Compensated orifice

4. Non-compensated orifice

| a. | 1 | c. | 1, 2, 3, and 4 |
| --- | --- | --- | --- |
| b. | 2 | d. | 1 and 2 |

ANS:	D

There are two types of flow restrictors: fixed-orifice and adjustable, multiple-orifice models. The adjustable, multiple-orifice flow restrictor uses a series of calibrated openings in a disk that can be adjusted to deliver different flows.

PTS:	1	REF:	Page 87

13.	The characteristic(s) of low-flow devices is (are):

1. fractional inspired oxygen (FIO2) from 0.22 to 0.60.

2. make use of a nasal cannula.

3. make use of a Venturi mask.

4. make use of a transtracheal catheter.

| a. | 1, 2, and 4 | c. | 1 |
| --- | --- | --- | --- |
| b. | 1 and 4 | d. | 1, 2, 3, and 4 |

ANS:	A

All of the choices are low-flow devices except for a Venturi mask. A Venturi mask is an air-entrainment device, all of which are considered to be high-flow devices. Therefore, Venturi masks deliver a precise FIO2.

PTS:	1	REF:	Page 98

14.	A patient with chronic bronchitis comes to your emergency department; she is discovered to have a pulse oximetry oxygen saturation (SpO2) of 78% with room air. On the basis of these data, which oxygen-delivery device would you use with the patient?

| a. | 6-L nasal cannula | c. | Simple mask |
| --- | --- | --- | --- |
| b. | Transtracheal catheter | d. | Venturi mask |

ANS:	D

A patient with chronic obstructive pulmonary disease requires a precise FIO2 because of carbon dioxide (CO2) retention issues. The only precise FIO2 device is a Venturi mask, which is considered a high-flow device.

PTS:	1	REF:	Pages 98-99

15.	High-flow devices include which of the following?

1. Nasal cannula

2. Venturi mask

3. Mask with reservoir bag

4. Oxygen hoods

| a. | 1 and 2 | c. | 3 and 4 |
| --- | --- | --- | --- |
| b. | 2 and 3 | d. | 2 and 4 |

ANS:	D

High-flow devices are also called fixed-performance devices because they provide oxygen at flow rates high enough to completely satisfy a patient’s inspiratory demands. These high-flow oxygen-delivery devices include air-entrainment masks (Venturi masks), incubators, oxygen tents, oxygen hoods, and high-volume aerosol devices.

PTS:	1	REF:	Page 90

16.	Which of the following are disadvantages of the nasal cannula?

1. Pressure necrosis

2. Drying nasal mucosa

3. FIO2 not precise

4. Claustrophobia

| a. | 1 and 4 | c. | 1, 2, and 3 |
| --- | --- | --- | --- |
| b. | 1 and 2 | d. | 2 and 3 |

ANS:	C

With straight prongs, oxygen flow is directed toward the superior aspects of the nasal cavity, thus promoting turbulent flow; with curved prongs, oxygen entering the nose is directed across the nasal turbinate, thus enhancing laminar flow as the gas flows through the nasal cavity. Thus, both types of prongs can cause the aforementioned disadvantages.

PTS:	1	REF:	Page 93

17.	If the air-entrainment port on a Venturi mask is occluded, which of the following will happen?

1. A decrease in flow

2. A decrease in FIO2

3. An increase in flow

4. An increase in FIO2

| a. | 1 and 2 | c. | 1 and 4 |
| --- | --- | --- | --- |
| b. | 1 and 3 | d. | None of the choices |

ANS:	C

Partial obstruction of oxygen flow downstream of the jet orifice or partial obstruction of the entrainment ports will decrease the amount of room air entrained, thus raising the FIO2 of the delivered gas.

PTS:	1	REF:	Page 98

18.	In theory, a nasal cannula can deliver what amount of FIO2 to a patient?

| a. | 0.24 to 0.60 | c. | 0.24 to 1.00 |
| --- | --- | --- | --- |
| b. | 0.24 to 0.36 | d. | 0.24 to 0.44 |

ANS:	D

The FIO2 levels delivered by a nasal cannula vary with the flow rate setting. Generally, the FIO2 increases by about 4% for each liter of flow increase beginning with 1 L/min delivering approximately 24% oxygen.

PTS:	1	REF:	Page 93, Table 4-1

19.	O2-induced hypoventilation is more likely to occur with which oxygen-delivery device?

| a. | Simple mask | c. | Venturi mask |
| --- | --- | --- | --- |
| b. | Nasal cannula | d. | Large-volume nebulizer |

ANS:	A

Excessive amounts of oxygen can cause a patient with chronically elevated CO2 levels to hypoventilate because the oxygen can knock out the patient’s hypoxic drive. The oxygen concentration delivered can vary significantly and may cause ventilatory depression in patients with chronically elevated partial pressure of arterial CO2 (PaCO2) levels.

PTS:	1	REF:	Page 89, Clinical Practice Guideline 4-1

20.	A nasal cannula that delivers 5 L/min translates to what approximate FIO2?

| a. | 36% | c. | 40% |
| --- | --- | --- | --- |
| b. | 28% | d. | 44% |

ANS:	C

The FIO2 increases approximately 4% for each liter of flow increase. If we begin with 20%, then 5 L/min  4 = 20. And 20 + 20% = 40%.

PTS:	1	REF:	Page 93, Table 4-1

21.	What liter flow is approximately 32% on a nasal cannula?

| a. | 2 | c. | 4 |
| --- | --- | --- | --- |
| b. | 3 | d. | 5 |

ANS:	B

For this calculation, zero flow is 20%. Therefore, 32% – 20% = 12%. If each liter represents 4% oxygen, then 12/4 = 3 L/min.

PTS:	1	REF:	Page 93, Table 4-1

22.	If a Venturi mask is set at 28% and the liter flow is set above the set liter flow prescribed on the device, what FIO2 is being delivered?

| a. | 28% |
| --- | --- |
| b. | 32% |
| c. | 100% |
| d. | 80% |

ANS:	A

The flow of oxygen to the nozzle is constant, so an obstruction will increase the FIO2 but not the flow.

PTS:	1	REF:	Page 99, Table 4-2

23.	Which of the following influences the FIO2 delivered to the patient by nasal

cannula?

1. Work of breathing

2. Tidal volume

3. Respiratory rate

4. Humidification

| a. | 4 | c. | 1 and 3 |
| --- | --- | --- | --- |
| b. | 1, 2, and 3 | d. | 1, 2, 3, and 4 |

ANS:	B

All physiologic aspects play a role in the amount of FIO2 delivered to a patient regardless of whether breathing is predominantly occurring through the nose or the mouth.

PTS:	1	REF:	Page 100

24.	Upon arrival to the emergency department, a patient is found to have an SpO2 of 75% on a nonrebreathing mask. You notice that the nonrebreathing mask is hooked to a bubble humidifier. The humidifier bottle is whistling and is expanded. What is your next action as the respiratory therapist?

| a. | Change the bubble humidifier; it’s defective. |
| --- | --- |
| b. | Check for kinks in tubing. |
| c. | Remove the nonrebreathing mask from the humidifier and run off of O2. |
| d. | Place a Venturi mask on the patient. |

ANS:	C

The humidity bottle is not designed for high flows. When it is set to a high flow, the humidity bottle expands and whistles; in turn there is a greater likelihood that the oxygen flowmeter will pop off from the pressure source.

PTS:	1	REF:	Page 94

25.	A patient with pulmonary fibrosis comes to the emergency department with an SpO2 of 87% and a respiratory rate of 28 breaths/min. The patient is unsure whether her transtracheal catheter is working properly; what is your first action?

| a. | Intubate her. | c. | Place a nasal cannula on her. |
| --- | --- | --- | --- |
| b. | Examine her for a mucus plug. | d. | Place a lavage catheter. |

ANS:	C

The therapist should think to place the patient on O2 to maintain her oxygenation status. Once the oxygen is running and the patient is breathing well, the therapist can examine the catheter for obstructions.

PTS:	1	REF:	Page 91

26.	The benefits of a transtracheal catheter include which of the following?

1. Lower flows are needed to obtain higher SpO2

2. The ability for the patient to eventually wean off of O2.

3. Improved self-image of the patient.

4. Oxygen cost savings of up to 60%.

| a. | 2 and 4 | c. | 1, 3, and 4 |
| --- | --- | --- | --- |
| b. | 1 and 3 | d. | 1, 2, 3, and 4 |

ANS:	C

The guiding principle for transtracheal oxygen is that the oxygen is delivered directly into the trachea, which provides the patient with adequate oxygen while reducing the amount of oxygen used. This reduces dilution with room air on inspiration. Consequently, lower oxygen flows from the source gas are required to achieve the desired arterial oxygen saturation (SaO2) with a cost savings of up to 60%.

PTS:	1	REF:	Page 94

27.	Complications of transtracheal oxygen catheters are:

| a. | subcutaneous emphysema. | c. | infection. |
| --- | --- | --- | --- |
| b. | mucus-plugging. | d. | all of the above. |

ANS:	D

These risks can be minimized with proper care, including saline instillation and periodic clearing of the catheter lumen with a guide wire or cleaning rod.

PTS:	1	REF:	Page 94

28.	Patients exposed to FIO2 greater than or equal to 50% are at risk for which of the following?

1. Chronic obstructive pulmonary disease

2. O2 toxicity

3. Permanent lung damage

4. Atelectasis

| a. | 1 and 2 | c. | 2, 3, and 4 |
| --- | --- | --- | --- |
| b. | 2 and 4 | d. | None of the choices |

ANS:	C

Oxygen is considered a type of medication. If a patient is exposed to large amounts for a long period of time, all of the aforementioned can occur except for chronic obstructive pulmonary disease, which is considered a disease process not a complication.

PTS:	1	REF:	Page 89, Clinical Practice Guideline 4-1

29.	In a pulse-demand oxygen-delivery system, the _____ replaces the flowmeter.

| a. | Pendant | c. | Pressure-relief valve |
| --- | --- | --- | --- |
| b. | Demand valve | d. | Flow valve |

ANS:	B

The demand valve connects directly to the oxygen source (50 psig), therefore replacing the flowmeter that is used with continuous-flow cannulas. Note that demand systems can function as pulsed or continuous-flow sources of oxygen. Settings allow the operator to select the equivalent of 1 to 5 L/min of oxygen flow from a conventional flowmeter.

PTS:	1	REF:	Page 96

30.	Types of reservoir cannulas include which of the following?

1. Mustache

2. Pulse demand

3. Pendant

4. High flow

| a. | 1 and 2 | c. | 1, 3, and 4 |
| --- | --- | --- | --- |
| b. | 2 and 3 | d. | 2, 3, and 4 |

ANS:	C

The two types of reservoir nasal cannulas are the mustache and the pendant. The high-flow nasal cannula can provide heated, humidified oxygen at flows varying between 5 and 40 L/min.

PTS:	1	REF:	Page 90

31.	Partial rebreathing masks deliver what FIO2 at what liter flow?

| a. | 0.35 to 0.50 at 4-5 L | c. | 0.50 to 0.80 at 6-8 L |
| --- | --- | --- | --- |
| b. | 0.40 to 0.60 at 6-8 L | d. | 0 50 to 1.00 at 12-15 L |

ANS:	B

A partial rebreathing mask can deliver an FIO2 of 0.40 to 0.60 with oxygen flows of 6 to 8 L/min. The actual percentage of oxygen delivered is also influenced by the patient’s ventilatory pattern.

PTS:	1	REF:	Page 97

32.	Which of the following are O2-conserving devices?

1. Transtracheal oxygen catheters

2. Reservoir cannulas

3. Pulse-demand oxygen-delivery systems

4. E size cylinder O2 tanks

| a. | 1, 2, and 4 | c. | 1, 3, and 4 |
| --- | --- | --- | --- |
| b. | 1, 2, and 3 | d. | 1, 2, 3, and 4 |

ANS:	B

Transtracheal oxygen catheters, reservoir cannulas, and pulse-demand oxygen-delivery systems are recent developments that have significantly improved the delivery of oxygen therapy, especially with regard to conserving oxygen supplies during long-term oxygen therapy.

PTS:	1	REF:	Page 90, Clinical Practice Guideline 4-2

33.	The transtracheal oxygen catheter is a guiding principle in which oxygen is delivered directly into the trachea and does which of the following?

| a. | Provides the patient with adequate oxygen |
| --- | --- |
| b. | Reduces the amount of oxygen used |
| c. | Provides an overall oxygen savings of 54% to 59% |
| d. | All of the above |

ANS:	D

That is, the direct delivery of oxygen into the trachea reduces dilution with room air on inspiration because the upper airways (the anatomic reservoir) are filled with oxygen.

PTS:	1	REF:	Page 94

34.	Which of the following are advantages of transtracheal catheters?

1. Increase patient mobility

2. Improve patient compliance because of an enhanced cosmetic appearance

3. Increase the amounts of nasal irritation

4. Heighten the risk for mucus plugging

| a. | 1, 2, and 3 | c. | 1 and 2 |
| --- | --- | --- | --- |
| b. | 2 and 3 | d. | 1, 2, 3, and 4 |

ANS:	C

Not only do transtracheal catheters produce an overall oxygen savings, they improve patient compliance with oxygen therapy because of cosmetic appearance. Transtracheal catheters are relatively inconspicuous. It is also found that patients with transtracheal catheters have increased mobility.

PTS:	1	REF:	Page 94

35.	The statement that is true concerning reservoir cannulas is which of the following?

| a. | The mustache cannula can hold about 40 mL of gas. |
| --- | --- |
| b. | The pendant cannula can reduce oxygen supply use by 75%. |
| c. | Patients prefer the mustache cannula to the pendant cannula. |
| d. | Pendant cannulas can be concealed by the patient’s clothing. |

ANS:	D

Studies indicate that mustache and pendant systems may reduce oxygen supply use by 50%. Many patients feel that mustache cannulas are heavier, larger, and more obvious. Pendant cannulas can be concealed by the patient’s clothing.

PTS:	1	REF:	Page 96

36.	The main goal of a pulse-demand oxygen-delivery system is which of the following?

1. To provide oxygen only during expiration

2. To provide oxygen only during inspiration

3. To deliver oxygen to the patient only after a sufficient inspiratory effort is made and the demand valve opens

4. To deliver oxygen at a preset flow rate and close during inspiration to conserve oxygen

| a. | 1 and 3 | c. | 2 and 3 |
| --- | --- | --- | --- |
| b. | 2 and 4 | d. | 1 and 4 |

ANS:	C

Pulse-demand oxygen-delivery systems deliver oxygen to the patient on demand only during inspiration. Oxygen is delivered to the patient only after a sufficient inspiratory effort is made and the demand valve opens. The demand valve closes during exhalation to conserve oxygen.

PTS:	1	REF:	Page 96

37.	A simple mask can deliver what percentage of FIO2 at how many liters/minute?

| a. | 0.24 to 0.44 at 1-6 L/min | c. | 0.40 to 0.60 at 10-15 L/min |
| --- | --- | --- | --- |
| b. | 0.35 to 0.50 at 5-10 L/min | d. | 0.50 to 0.80 at 5-10 L/min |

ANS:	B

The simple mask can deliver a higher concentration than a nasal cannula because of the volume of the mask itself. That volume serves as an extension of the anatomic reservoir. The simple mask is generally said to be able to deliver 0.35 to 0.50 at oxygen flows of 5 to 10 L/min.

PTS:	1	REF:	Page 93

38.	The FIO2 that is actually delivered to the patient through an oxygen mask depends on which of the following?

1. Flow of the oxygen to the mask

2. Presence of a reservoir bag

3. Humidity of the gas

4. Density of the gas

| a. | 1 and 2 | c. | 1 and 4 |
| --- | --- | --- | --- |
| b. | 2 and 3 | d. | 3 and 4 |

ANS:	A

The flow of oxygen to a reservoir mask determines the FIO2, as do the patient’s breathing pattern and the presence or absence of a reservoir bag. Those masks with a reservoir bag will have a higher FIO2 than those without a reservoir bag.

PTS:	1	REF:	Pages 93-94

39.	Which of the following devices could cause CO2 retention?

1. Simple mask

2. Aerosol face mask

3. Nonrebreathing mask

4. Air-entrainment mask

| a. | 1 and 3 | c. | 2 and 3 |
| --- | --- | --- | --- |
| b. | 1 and 2 | d. | 3 and 4 |

ANS:	A

Carbon dioxide rebreathing can occur if the oxygen flow to a simple or nonrebreathing mask is not sufficient to wash out the patient’s exhaled gases. The flow of oxygen to a reservoir mask determines the FIO2, as do the patient’s breathing pattern and the presence or absence of a reservoir bag. Those masks with a reservoir bag will have a higher FIO2 than those without a reservoir bag.

PTS:	1	REF:	Page 97

40.	In a typical adult partial rebreathing mask, the reservoir bag has an approximate volume capacity of _____ mL?

| a. | 150-300 | c. | 400-600 |
| --- | --- | --- | --- |
| b. | 300-500 | d. | 500-700 |

ANS:	B

A typical partial rebreathing mask has a reservoir bag with a volume capacity of 300 to 500 mL.

PTS:	1	REF:	Page 97

41.	Which of the following O2 devices has 2 one-way valves and the second valve is at the exhalation ports, which prevents room air from entering the mask during inhalation?

| a. | Nonrebreathing mask | c. | Aerosol tracheostomy collar mask |
| --- | --- | --- | --- |
| b. | Partial rebreathing mask | d. | Simple mask |

ANS:	A

The first set of valves is a one-way valve located between the reservoir bag and the base of the mask. This valve allows gas flow to enter the mask from the reservoir bag when the patient inhales and prevents gas flow from the mask flowing back into the reservoir bag during the patient’s exhalation. The second set of valves is at the exhalation ports. The one-way valves placed there prevent room air from entering the mask during inhalation. They also allow the patient’s exhaled gases to exit the mask during exhalation.

PTS:	1	REF:	Pages 97-98

42.	Theoretically, a nonrebreathing mask can deliver 100% oxygen, assuming that the mask fits appropriately on the patient’s face and is the only source of gas being inhaled by the patient. However, in actual practice, the disposable nonrebreathing mask can deliver what percentage of FIO2?

| a. | 0.30-0.50 | c. | 0.60-0.80 |
| --- | --- | --- | --- |
| b. | 0.40-0.60 | d. | 0.24-0.50 |

ANS:	C

The discrepancy between disposable nonrebreathing masks and the original Boothby-Lovelace-Bulbulian masks is primarily related to the fact that manufacturers usually supply disposable masks with one of the exhalation valves removed.

PTS:	1	REF:	Page 98

43.	The concentration of oxygen delivered to the patient with an air-entrainment mask depends on which of the following?

1. Flow of the oxygen exiting the jet nozzle

2. Size of the jet nozzle outlet

3. Size of the entrainment port

4. Size of the Venturi mask

| a. | 1, 2, and 3 | c. | 1, 2, and 4 |
| --- | --- | --- | --- |
| b. | 2, 3, and 4 | d. | 3 and 4 |

ANS:	A

Oxygen flowing through the jet nozzle of an air-entrainment device “drags” in room air through the entrainment ports as a result of viscous, shearing forces between the gas exiting the jet nozzle outlet and the surrounding ambient air. The concentration of oxygen delivered to the patient, therefore, depends on the flow of oxygen exiting the jet nozzle, the size of the jet nozzle outlet, and the size of the entrainment ports.

PTS:	1	REF:	Page 98

44.	The total flow for a 35% air-entrainment mask running at 10 L/min is _____ L/min.

| a. | 43 | c. | 60 |
| --- | --- | --- | --- |
| b. | 53 | d. | 65 |

ANS:	B

100% O2 – 35% FIO2 = 65%. 20% air – 35% FIO2 = 15. Ratio: 15:65 = 1:4.3. For every 1 L/min of O2, there is 4.3 L/min of air being entrained. 1 + 4.3 = 5.3  10 = 53 L/min total flow.

PTS:	1	REF:	Page 99, Figure 4-21

45.	How many liters/minute of air is being entrained for a 35% air-entrainment mask running at 10 L/min?

| a. | 43 L/min | c. | 60 L/min |
| --- | --- | --- | --- |
| b. | 53 L/min | d. | 65 L/min |

ANS:	A

Refer to question 44. Air: 4.3  10 L/min = 43 L/min of air being entrained.

PTS:	1	REF:	Page 99, Figure 4-21

46.	The oxygen/air ratio for the 35% air-entrainment mask is:

| a. | 1:4.3. | c. | 1:6.0. |
| --- | --- | --- | --- |
| b. | 1:5.3. | d. | 1:6.5. |

ANS:	A

Refer to question 44.

PTS:	1	REF:	Page 99, Figure 4-21

47.	The total flow for a 50% air-entrainment mask running at 8 L/min is _____ L/min.

| a. | 18 | c. | 25 |
| --- | --- | --- | --- |
| b. | 22 | d. | 30 |

ANS:	B

100% O2 – 50% = 50%. 20% air – 50% FIO2 = 30. 30:50 = 1:1.7. For every 1 L/min of O2, there is 1.7 L/min of air being entrained. 1 + 1.7 = 2.7  8 = 21.6 L/min total flow.

PTS:	1	REF:	Page 99, Figure 4-21

48.	How many liters/minute of air is being entrained for a 50% air-entrainment mask running at 8 L/min?

| a. | 10 | c. | 14 |
| --- | --- | --- | --- |
| b. | 12 | d. | 20 |

ANS:	C

Refer to question 47. 1.7  8 = 13.6 L/min of air being entrained.

PTS:	1	REF:	Page 99, Figure 4-21

49.	A 22-year-old with asthma comes to the emergency department in respiratory distress. The patient has acute shortness of breath (SOB); is febrile, tachypneic, and tachycardic; and has an SpO2 of 88% on a nasal cannula that delivers 2 L/min. Which of the following actions should the respiratory therapist take at this time?

| a. | Teach the patient to breathe through pursed lips. |
| --- | --- |
| b. | Place a simple mask on the patient. |
| c. | Remove the cannula and place a 0.35 air-entrainment mask on the patient. |
| d. | Place a high-flow cannula at 15 L/min on the patient. |

ANS:	C

A patient with asthma benefits from a precise preset FIO2. With a simple mask, there is a greater likelihood of the patient retaining CO2, which is not helpful for those with asthma. Breathing through pursed lips is beneficial to people with asthma; however, it is not the first course of action. Finally, a high-flow cannula at 15 L/min is too much FIO2 for the patient at this time.

PTS:	1	REF:	Page 98

50.	Which of the following O2 devices is appropriate for patients with chronic obstructive pulmonary disease who are hypoxemic?

| a. | Simple mask | c. | Partial rebreathing mask |
| --- | --- | --- | --- |
| b. | Nasal cannula | d. | Air-entrainment mask |

ANS:	D

Air-entrainment masks are excellent for providing oxygen therapy to hypoxemic patients with chronic obstructive pulmonary disease who typically require a fixed FIO2 between 0.24 and 0.35. The total flow of gas delivered (oxygen plus air) by such masks for lower FIO2 levels is usually sufficient to meet the peak inspiratory flow requirements for these patients.

PTS:	1	REF:	Page 98

51.	With oxygen hoods, FIO2 must be measured by using what type of device?

| a. | Incubator | c. | Heater |
| --- | --- | --- | --- |
| b. | Pulse oximeter | d. | Analyzer |

ANS:	D

Oxygen hoods must be measured intermittently or monitored continuously with an oxygen analyzer.

PTS:	1	REF:	Page 100

52.	Current incubators are designed to allow all of the following, except:

| a. | Variable control of environmental temperature |
| --- | --- |
| b. | Variable control of environmental humidity |
| c. | Control of FIO2 concentration |
| d. | Control of barometric pressure |

ANS:	D

The temperature and humidity of the gas within the incubator are regulated by a servo-controlled mechanism connected to a fan that circulates environmental gas over heating coils and a blow-by humidifier. Supplemental oxygen can be provided by connecting a heated humidifier directly to the incubator.

PTS:	1	REF:	Page 100

53.	An oxygen _____ compresses air and oxygen from a high-pressure source into a chamber where the pressure of the two gases is equalized, usually at 50 psi.

| a. | blender | c. | hood |
| --- | --- | --- | --- |
| b. | tent | d. | delivery devices |

ANS:	A

An oxygen blender is attached to both a high-pressure compressed air source and a high-pressure oxygen source of equal pressure (approximately 50 psig).

PTS:	1	REF:	Page 101

54.	It is important to filter the gas before it enters the blender housing; otherwise, a malfunction will occur. Which of the following can cause a malfunction?

1. Moisture

2. Particulate matter

3. High amounts of FIO2

4. Low amounts of FIO2

| a. | 1 and 3 | c. | 1 and 4 |
| --- | --- | --- | --- |
| b. | 1 and 2 | d. | 2 and 3 |

ANS:	B

Because moisture and particulate matter introduced into the blender by the source gases can cause the blender to malfunction, it is important to filter the gas before it enters the blender housing.

PTS:	1	REF:	Pages 101-102

55.	_____ therapy exposes patients to a pressure greater than atmospheric pressure while they breathe 100% oxygen, either continuously or intermittently.

| a. | Aerosol oxygen | c. | Bronchial hygiene |
| --- | --- | --- | --- |
| b. | Hyperbaric oxygen | d. | Nitric oxide |

ANS:	B

Hyperbaric oxygen therapy exposes patients to a pressure greater than atmospheric pressure while they breathe 100% oxygen, either continuously or intermittently.

PTS:	1	REF:	Page 102

56.	What types of patients are treated with hyperbaric oxygen therapy?

1. Deep sea diving patients with decompression sickness

2. Those with thermal injuries

3. Patients with carbon monoxide poisoning

4. Those with refractory osteomyelitis

| a. | 1, 2, and 3 | c. | 2, 3, and 4 |
| --- | --- | --- | --- |
| b. | 1, 2, and 4 | d. | 1, 2, 3, and 4 |

ANS:	D

Decompression sickness arises from the generation of nitrogen bubbles in the vascular system and tissues in volumes sufficient to interfere with the function of an organ. The cause of gas bubble formation is the rapid decompression during ascent from diving when the speed of decompression exceeds the ability of the gas-saturated tissues to vent the gases by simple diffusion. For this type of patient, hyperbaric oxygen will reduce the size of the bubbles. For thermal injuries, hyperbaric oxygen limits the progression of the burn injury, reduces swelling, and may reduce the need for surgical intervention. For CO poisoning, hyperbaric oxygen promotes the dissociation of carboxyhemoglobin more than just oxygen alone. For refractory osteomyelitis, hyperbaric oxygen provides periodic elevation of bone and tissue oxygen tension from hypoxic levels to normal or supranormal levels, which promotes collagen production by fibroblasts.

PTS:	1	REF:	Page 105, Box 4-1

57.	Exposure to elevated barometric pressure during hyperbaric oxygen therapy can directly affect a number of physiologic parameters, including which of the following?

1. Arterial and alveolar partial pressure for oxygen

2. Temperature of the gases being breathed

3. Work of breathing

4. Lung volumes

| a. | 1 and 2 | c. | 2, 3, and 4 |
| --- | --- | --- | --- |
| b. | 1 and 4 | d. | 1, 2, 3, and 4 |

ANS:	D

The effect of increased ambient pressure on the partial pressure of alveolar oxygen (PAO2) can be explained by Dalton’s law, which states that the total pressure of a gas mixture, such as air, equals the sum of the partial pressures of each of the constituent gases in the mixture. According to Gay-Lussac’s law, if the volume of a gas remains constant, there is a direct relationship between the absolute pressure of a gas and its temperature. It is reasonable to suggest that if the volume of a hyperbaric chamber remains constant, then increasing the pressure would raise the temperature inside of the chamber.

PTS:	1	REF:	Page 102

58.	______ law states that if the temperature of a gas remains constant, the volume of a gas is inversely related to its pressure.

| a. | Henry’s | c. | Dalton’s |
| --- | --- | --- | --- |
| b. | Boyle’s | d. | Laplace’s |

ANS:	B

Boyle's law states that if the temperature of a gas remains constant, the volume of a gas is inversely related to its pressure. Therefore, as pressure is exerted on a container, the gas volume within the container decreases.

PTS:	1	REF:	Page 102

59.	Dalton’s law states that:

| a. | The total pressure of a gas mixture equals the sum of the partial pressure of each constituent gas in the mixture. |
| --- | --- |
| b. | The pressure within a liquid sphere is influenced by the surface tension forces of the liquid and the size of the sphere. |
| c. | There is a direct relationship between the absolute pressure of a gas and its temperature. |
| d. | The volume of a gas is inversely related to its pressure. |

ANS:	A

Dalton’s law states that the sum of the partial pressures of a gas mixture equals the total pressure of the system. Laplace’s law states that the pressure within a liquid sphere is influenced by the surface tension forces of the liquid and the size of the sphere. Gay-Lussac demonstrated that if the volume of a gas is held constant, the gas pressure rises as the absolute temperature of the gas increases. Boyle’s law states that when the temperature is held constant, the volume that a gas occupies is inversely proportional to the absolute pressure exerted on it.

PTS:	1	REF:	Page 102

60.	A patient is receiving 0.30 FIO2 with a barometric pressure at 750 mm Hg. The partial pressure of carbon dioxide (PCO2) for this patient is 70 mm Hg; the partial pressure of arterial oxygen (PaO2) is 60 mm Hg. The PAO2 for this patient is _____ mm Hg.

| a. | 100.2 | c. | 160.0 |
| --- | --- | --- | --- |
| b. | 123.4 | d. | 181.7 |

ANS:	B

Formula: PAO2 = (Barometric pressure [PB] – pressure of water vapor [PH2O]) FIO2 – PaCO2 (1.25).

PAO2 = (750 mm Hg – 47) 0.30 – 70 (1.25).

PAO2 = 210.9 – 87.5.

PAO2 = 123.4 mm Hg.

PTS:	1	REF:	Page 103

61.	A patient is receiving 0.60 FIO2 with a barometric pressure of 755 mm Hg. PCO2 is 56 mm Hg and PaO2 65 mm Hg. The PAO2 for this patient is _____ mm Hg.

| a. | 265.6 | c. | 354.8 |
| --- | --- | --- | --- |
| b. | 300.9 | d. | 404.0 |

ANS:	C

Formula: PAO2 = (PB – PH2O) FIO2 – PaCO2 (1.25).

PAO2 = (755 – 47) 0.60 – 56 (1.25).

PAO2 = 424.8 – 70.

PAO2 = 354.8 mm Hg.

PTS:	1	REF:	Page 103

62.	A patient is receiving 0.80 FIO2 with a barometric pressure at 760 mm Hg. PCO2 is at 68; PaO2 is 55. The PAO2 for this patient is _____ mm Hg.

| a. | 395.6 | c. | 510.9 |
| --- | --- | --- | --- |
| b. | 485.4 | d. | 550.0 |

ANS:	B

Formula: PAO2 = (PB – PH2O) FIO2 – PaCO2 (1.25).

PAO2 = (760 – 47) 0.80 – 68 (1.25).

PAO2 = 570.4 – 85.

PAO2 = 485.4 mm Hg.

PTS:	1	REF:	Page 103

63.	As the barometric pressure increases, there is an increase in the density of the gas being breathed. The increase in gas density in turn causes a(n):

| a. | Decrease in work of breathing | c. | Increase in work of breathing |
| --- | --- | --- | --- |
| b. | Increase in FIO2 | d. | Decrease in FIO2 |

ANS:	C

The increase in gas density results in increased work of breathing, which is not noticeable and can easily be accommodated in healthy subjects. In patients with reduced lung reserves, however, this increased work can present problems and necessitate ventilatory support.

PTS:	1	REF:	Page 103

64.	Intermittent air-breathing is used in hyperbaric chambers to prevent which of the following?

| a. | Hyperventilation | c. | Hypoventilation |
| --- | --- | --- | --- |
| b. | Oxygen toxicity | d. | Hypertension |

ANS:	B

Hyperbaric oxygenation is achieved by having the patient breathe oxygen by mask or through a specially designed hood while being exposed to elevated barometric pressures in the compressed-air chamber. Treatment schedules are tailored to the specific needs of the patient. Generally, patients are placed on schedules in which intermittent air-breathing periods of 5 minutes or more are programmed approximately every 20 minutes to prevent oxygen toxicity to the patient.

PTS:	1	REF:	Page 102

65.	Which types of devices/techniques should be used to monitor the oxygenation status of patients undergoing hyperbaric oxygen therapy?

1. Transcutaneous monitoring

2. Arterial blood gas monitoring

3. End-tidal CO2 monitoring

4. Spirometer

| a. | 1 and 2 | c. | 2 and 3 |
| --- | --- | --- | --- |
| b. | 1 and 4 | d. | 3 and 4 |

ANS:	A

Transcutaneous monitoring has proven to be valuable in the assessment of the overall oxygenation status of a patient undergoing hyperbaric therapy. Arterial blood gas monitoring can provide information on the oxygenation status of patients receiving hyperbaric therapy and can also give an indication of their ventilatory status.

PTS:	1	REF:	Page 104

66.	Contraindications for hyperbaric oxygen therapy consist of all of the following, except:

| a. | Pneumothorax | c. | Upper respiratory infection |
| --- | --- | --- | --- |
| b. | Obstructive bronchial diseases | d. | Skin grafting |

ANS:	D

If pneumothorax occurs during hyperbaric treatment, chest tubes should be immediately inserted; failure to treat pneumothorax can have dire consequences. Gas-trapping can result in barotraumas for those with obstructive bronchial diseases. Similarly, patients who have upper respiratory infections and nasal congestion are usually unable to clear their ears during compression and decompression and thus are prone to eardrum rupture during treatment.

PTS:	1	REF:	Page 105, Box 4-2

67.	Which type of therapy has been used to successfully treat persistent pulmonary hypertension in newborns?

| a. | Helium-oxygen (heliox) | c. | Bronchial hygiene |
| --- | --- | --- | --- |
| b. | Nitric oxide | d. | Carbon dioxide–nitrogen (carbogen) |

ANS:	B

Nitric oxide has been shown to be a potent pulmonary vasodilator.

PTS:	1	REF:	Page 105

68.	The therapeutic dose of nitric oxide is _____ ppm.

| a. | 2-80 | c. | 40-150 |
| --- | --- | --- | --- |
| b. | 25-115 | d. | 30-60 |

ANS:	A

The therapeutic dose of nitric oxide is 2 to 80 ppm.

PTS:	1	REF:	Page 105

69.	Heliox has been used on a limited basis to treat patients with which of the following?

| a. | Hyperventilation | c. | Airway obstructions |
| --- | --- | --- | --- |
| b. | Thermal burns | d. | Upper respiratory infections |

ANS:	C

Specifically, heliox has been used to manage asthmatic patients with acute respiratory failure, to treat postextubation stridor in pediatric trauma patients, as an adjunct in the treatment of pediatric patients with refractory croup, to administer anesthetic gases to patients via small-diameter endotracheal tubes, and to provide ventilatory support for patients with severe airway obstruction caused by chronic bronchitis and emphysema.

PTS:	1	REF:	Page 106

70.	An 18-year-old with severe asthma is brought into the emergency department on an nonrebreathing mask; he is in acute respiratory distress. He has been given multiple rounds of bronchodilators and corticosteroids and is still in distress, with an SpO2 of 85% on the nonrebreathing mask. What therapy should the respiratory care practitioner suggest to the physician at this time?

| a. | Nitric oxide therapy | c. | Carbogen therapy |
| --- | --- | --- | --- |
| b. | Heliox therapy | d. | Hyperbaric therapy |

ANS:	B

Heliox mixtures have been used on a limited basis to treat patients with airway obstruction, specifically in the management of asthmatic patients with acute respiratory failure.

PTS:	1	REF:	Page 106

71.	A patient is placed on a 70:30 heliox mixture at a flow of 8 L/min. The actual flow rate going to the patient is ____ L/min.

| a. | 11.2 | c. | 14.4 |
| --- | --- | --- | --- |
| b. | 12.8 | d. | 16.2 |

ANS:	B

Conversion for 70:30 = 1.6  8 L/min = 12.8 L/min.

PTS:	1	REF:	Page 107, Table 4-3

72.	A pediatric patient with refractory croup is placed on heliox 80:20 at a flow of 10 L/min. The actual flow going to the patient is _____ L/min.

| a. | 14 | c. | 18 |
| --- | --- | --- | --- |
| b. | 16 | d. | 20 |

ANS:	C

Conversion for 80:20 = 1.8  10 L/min = 18 L/min.

PTS:	1	REF:	Page 107, Table 4-3

73.	The benefits of breathing heliox are related to its lower density in comparison with pure oxygen or air. The lower density does which of the following?

1. Promotes laminar flow

2. Increases the amount of turbulent flow

3. Reduces the amount of turbulent flow

4. Decreases the laminar flow

| a. | 1 and 2 | c. | 2 and 4 |
| --- | --- | --- | --- |
| b. | 1 and 3 | d. | 3 and 4 |

ANS:	B

This relationship is important to remember when administering heliox because the actual flow rate of gas delivered will be greater than the set flow.

PTS:	1	REF:	Page 106

74.	What type of oxygen device is used to administer heliox to a non-intubated patient?

| a. | Nasal cannula | c. | Simple mask |
| --- | --- | --- | --- |
| b. | High-flow cannula | d. | Nonrebreathing mask |

ANS:	D

For non-intubated patients, a well-fitted nonrebreathing mask attached to a reservoir bag should be used. The flow rate of gas should be high enough to prevent the reservoir bag from collapsing during inspiration. Nasal cannulas are ineffective for delivering heliox because of leakage; instead, nonrebreathing masks are the most effective.

PTS:	1	REF:	Page 106

75.	A 55-year-old man comes to the emergency department on a nonrebreathing mask with an SpO2 of 88%. The patient was found unconscious in a running car in the garage. Carbon monoxide poisoning is a concern at this time. What should the respiratory care practitioner recommend to the physician at this time?

| a. | Place a nasal cannula at 6 L/min on the patient. |
| --- | --- |
| b. | Begin the patient on heliox therapy 80%:20%. |
| c. | Place the patient on carbogen therapy 5%:95%. |
| d. | Set the patient up on an air-entrainment device at 0.50 FIO2. |

ANS:	C

Carbogen is used to treat hiccups and carbon monoxide poisoning; furthermore, it is used as a stimulant/depressant of ventilation and to prevent the complete washout of CO2 during cardiopulmonary bypass.

PTS:	1	REF:	Page 107, Table 4-3

76.	To prevent an adverse reaction, it is essential to monitor which of the following when administering carbogen to a patient?

1. Heart rate

2. Blood pressure

3. Respirations

4. The patient’s mental status

| a. | 1, 2, and 3 | c. | 2, 3, and 4 |
| --- | --- | --- | --- |
| b. | 1, 2, and 4 | d. | 1, 2, 3, and 4 |

ANS:	D

Pulse, arterial blood pressure, and minute volume normally increase as the patient breathes carbogen, but the rapidity and level of these changes depend on the concentration of the mixture.

PTS:	1	REF:	Page 107, Table 4-3

77.	According to the American Association for Respiratory Care (AARC) Clinical Practice Guideline (2002 Revision & Update), which of the following are considered precautions or complications of oxygen therapy?

1. A PaO2 >60 mm Hg in spontaneously breathing patients with chronic PaCO2 levels

2. With FIO2 >0.50 and the risk of oxygen toxicity or absorption atelectasis occurring

3. During laser bronchoscopy and high levels of oxygen

4. Acute myocardial infarction

| a. | 1, 2, and 4 | c. | 2, 3, and 4 |
| --- | --- | --- | --- |
| b. | 1, 2, and 3 | d. | 1, 2, 3, and 4 |

ANS:	B

See the AARC Clinical Practice Guideline: Oxygen Therapy for Adults in the Acute Care Facility—2002 Revision and Update, Clinical Practice Guideline 4-1.

PTS:	1	REF:	Page 89, Clinical Practice Guideline 4-1

78.	All oxygen-delivery systems should be checked how often according to the AARC Clinical Practice Guideline (2002 Revision & Update)?

| a. | At least once a day | c. | Every 48 hours |
| --- | --- | --- | --- |
| b. | Twice a week | d. | Once a month |

ANS:	A

See the AARC Clinical Practice Guideline: Oxygen Therapy for Adults in the Acute Care Facility—2002 Revision and Update, Clinical Practice Guideline 4-1.

PTS:	1	REF:	Page 89, Clinical Practice Guideline 4-1

79.	What criteria must a patient meet for home oxygen therapy, according to AARC Clinical Practice Guidelines (2002 Revision)?

1. PaO2 <55 mm Hg

2. SpO2 <88% room air at rest

3. SpO2 <88% room air with exertion

4. PaO2 >60 mm Hg

| a. | 1, 2, and 4 | c. | 1, 2, and 3 |
| --- | --- | --- | --- |
| b. | 2, 3, and 4 | d. | 2 and 3 |

ANS:	C

See the AARC Clinical Practice Guideline: Oxygen Therapy for Adults in the Acute Care Facility—2002 Revision and Update, Clinical Practice Guideline 4-1.

PTS:	1	REF:	Page 89, Clinical Practice Guideline 4-1

80.	Retinopathy of prematurity often occurs in neonates because of which of the following?

| a. | Increased amounts of FIO2 |
| --- | --- |
| b. | Low amounts of FIO2 |
| c. | Improper placement of the oxygen-delivery device |
| d. | Hyperventilation |

ANS:	A

The origin of retinopathy of prematurity, especially with regard to the role of oxygen, is controversial. Nonetheless, care should be taken when supplemental oxygen is provided to preterm infants (<37 weeks’ gestation). It is suggested that oxygen supplementation should not result in a PaO2 >80 mm Hg.

PTS:	1	REF:	Page 91, Clinical Practice Guideline 4-3

81.	For a neonate on an oxygen hood, where is the best placement of an oxygen analyzer to measure the correct FIO2?

| a. | Above the patient’s head | c. | On the side near either arm |
| --- | --- | --- | --- |
| b. | At the nose and mouth | d. | The rear of the hood |

ANS:	B

Neonates are obligate nose-breathers, so the mouth and nose are the most accurate location, because it is the area closest to the nares.

PTS:	1	REF:	Page 93

82.	What is the percentage of FIO2 for a patient on 3 L/min delivered through a nasal cannula?

| a. | 32% | c. | 44% |
| --- | --- | --- | --- |
| b. | 40% | d. | 50% |

ANS:	A

3 L/min  4 = 12 L/min + 20 = 32% FIO2. The flow is multiplied by 4 because for every 1 L/min of oxygen flowing to a nasal cannula the FIO2 increases by 4%. The 20 is the number that the flow multiplied by 4 is added to, in order to get the approximate FIO2.

PTS:	1	REF:	Page 93, Table 4-1

83.	Which of the following are disadvantages of delivering oxygen by reservoir masks?

1. Rebreathing CO2

2. The patient cannot eat

3. Increased risk of aspirations

4. Claustrophobia/discomfort

| a. | 2, 3, and 3 | c. | 1, 3, and 4 |
| --- | --- | --- | --- |
| b. | 2, 3, and 4 | d. | 1, 2, 3, and 4 |

ANS:	D

Oxygen masks are confining and may not be well-tolerated by some patients. Furthermore, they must be removed during eating, drinking, and facial and airway care. Patients often report that these oxygen masks cause skin irritation, especially when they are tightly fitted. Finally, both aspiration of vomitus and rebreathing of CO2 are more likely when a mask is being used.

PTS:	1	REF:	Pages 96-97

84.	Which of the following are advantages of nasal cannula use?

1. They are cost-effective.

2. They deliver precise FIO2.

3. The patient can eat, speak, and drink with the cannula in place.

4. Pressure necrosis will not occur.

| a. | 1 and 2 | c. | 2 and 3 |
| --- | --- | --- | --- |
| b. | 1 and 3 | d. | 3 and 4 |

ANS:	B

Nasal cannulas are inexpensive and allow the patient to eat, drink, and speak while receiving oxygen. Nasal cannulas cannot deliver precise oxygen concentrations, and pressure necrosis is possible because of the pressure points where the tubing holds the cannula in place touch the patient’s face and ears.

PTS:	1	REF:	Page 97

85.	Which of the following are the potential harmful effects of excessive amounts of oxygen therapy?

1. Persistent pulmonary hypertension

2. Oxygen-induced hypoventilation

3. Absorption atelectasis

4. Retinopathy or prematurity

| a. | 1, 2, and 3 | c. | 2, 3, and 4 |
| --- | --- | --- | --- |
| b. | 1, 2, and 4 | d. | 1, 2, 3, and 4 |

ANS:	C

Oxygen is considered a type of medication. If exposed to large amounts of oxygen for a long period of time, a patient can experience lung damage because oxygen can be toxic if given in high doses.

PTS:	1

REF:	Page 89, Clinical Practice Guideline 4-1 | Page 91, Clinical Practice Guideline 4-3

86.	When a closed Thorpe tube flowmeter is attached to a 50-psi oxygen source, the indicator float jumps up then quickly falls to zero. This flowmeter is:

| a. | broken. | c. | compensated for back pressure. |
| --- | --- | --- | --- |
| b. | a Bourdon gauge flowmeter. | d. | uncompensated for back pressure. |

ANS:	C

A flowmeter can be determined to be pressure-compensated if the following test is performed: With the needle valve closed, the flowmeter is plugged into a high-pressure gas source. If the float in the indicator tube jumps and then falls to zero, the flowmeter is pressure-compensated. The float movement occurs because the source gas must pass through the indicator tube before it reaches the needle valve.

PTS:	1	REF:	Page 86

87.	What type of oxygen-delivery device does the following image illustrate?

| a. | Simple mask | c. | Air-entrainment mask |
| --- | --- | --- | --- |
| b. | Nonrebreathing mask | d. | Aerosol mask |

ANS:	B

See also Figure 4-19.

PTS:	1	REF:	Page 97

88.	What type of oxygen-delivery device does the following image illustrate?

| a. | Simple mask | c. | Air-entrainment mask |
| --- | --- | --- | --- |
| b. | Nonrebreathing mask | d. | Partial rebreathing mask |

ANS:	C

See Figure 4-20.

PTS:	1	REF:	Page 98, Figure 4-20

89.	______ law states that the degree to which a gas enters into a physical solution in body fluids is directly proportional to the partial pressure of gas to which the fluid is exposed?

| a. | Dalton’s | c. | Gay-Lussac’s |
| --- | --- | --- | --- |
| b. | Henry’s | d. | Boyle’s |

ANS:	B

Henry’s law is used to explain the changes in the PaO2 that occur with exposure to elevated ambient pressure. The law states that the degree to which a gas enters into physical solution in body fluids is directly proportional to the partial pressure of gas to which the fluid is exposed.

PTS:	1	REF:	Page 103

90.	A 6-year-old patient with asthma has a sudden onset of SOB and wheezing. He presents to the emergency department with a heart rate of 82 beats/min, respiratory rate of 28 breaths/minute, and an SpO2 of 87% on room air. Which of the following actions should the respiratory care practitioner take at this time?

1. Place the patient on oxygen therapy.

2. Intubate the patient.

3. Administer a bronchodilator.

4. Administer nitric oxide therapy.

| a. | 1 and 3 | c. | 2 and 4 |
| --- | --- | --- | --- |
| b. | 1 and 4 | d. | 3 and 4 |

ANS:	A

Immediate relief for the patient at this time can be provided by a bronchodilator to help open up the airways and by oxygen to improve the patient’s oxygenation status. Intubation and nitric oxide therapy can both be administered to people with asthma; however, these approaches are taken when the others do not work for the patient and depend on the severity of the patient’s condition.

PTS:	1	REF:	Page 89

91.	A 60-year-old man with a 90 pack-years smoking history arrives in the emergency department on a nonrebreathing mask with an SpO2 of 88%. The arterial blood gas results at this time indicate a pH of 7.38, PCO2 of 58, and partial pressure of oxygen (PO2) of 54. Which action should the respiratory therapist take at this time?

| a. | Keep the nonrebreathing mask on the patient. |
| --- | --- |
| b. | Administer an 80:20 heliox concentration. |
| c. | Place a nasal cannula set at 6 L/min. |
| d. | Place an air-entrainment mask on the patient that delivers 0.40 FIO2. |

ANS:	D

An air-entrainment device is the best delivery device for a patient with chronic obstructive pulmonary disease because of its ability to provide a precise FIO2. The other choices have a greater likelihood of causing CO2 retention and knocking out the patient’s hypoxic drive.

PTS:	1	REF:	Page 98

92.	A home care patient calls to inform the respiratory therapist that his transtracheal catheter accidentally fell out last night and that he was unable to reinsert it. The patient should be told to:

| a. | insert a dilating or stenting device. |
| --- | --- |
| b. | continue attempts to reinsert the catheter. |
| c. | use a nasal cannula and call his physician as soon as possible. |
| d. | use a nasal cannula until another transtracheal catheter is delivered to the home. |

ANS:	D

Without the transtracheal catheter in place, this patient is not receiving any supplemental oxygen. To avoid hypoxemia, the patient must use a back-up delivery device. The most appropriate device in this situation would be a nasal cannula. The patient must call the physician for follow-up immediately.

PTS:	1	REF:	Page 95

93.	The partial pressure of alveolar oxygen (PAO2) when the atmospheric pressure is 3 atmospheric pressure (atm), the PACO2 is 40 mm Hg, and the FIO2 is 0.80 is _____ mm Hg.

| a. | 2280 | c. | 1736 |
| --- | --- | --- | --- |
| b. | 1784 | d. | 1500 |

ANS:	C

PAO2 = (Pbar – PH2O) FIO2 – PaCO2 ÷ 0.8; Pbar = 3 atm  760 = 2280 mm Hg.

PAO2 = (2280 – 47)0.8 – 40 ÷ 0.8.

PAO2 = 1786.4 – 50 = 1736.4 mm Hg.

PTS:	1	REF:	Page 103

94.	A 45-year-old man had a deep sea diving accident recently and developed air emboli. Which type of therapy should the therapist recommend at this time?

| a. | Nitric oxide therapy | c. | Hyperbaric therapy |
| --- | --- | --- | --- |
| b. | Heliox therapy | d. | Aerosol therapy |

ANS:	C

Hyperbaric oxygen therapy exposes patients to a pressure greater than atmospheric pressure while they breathe 100% oxygen either continuously or intermittently. The effects on lung volume can be explained by Boyle’s law, which states that if the temperature of a gas remains constant, the volume of a gas is inversely related to its pressure. That is, as the pressure exerted on the container increase, the gas volume decreases. Thus when a person is exposed to elevated pressures, the gas volume contained in any body cavity tends to be compressed.

PTS:	1	REF:	Page 105, Box 4-1

95.	A patient has a 0.40 air-entrainment mask on and is not tolerating it well at this time because of claustrophobia. A nasal cannula is placed on the patient. What liter flow should be set to match the appropriate FIO2?

| a. | 2 | c. | 4 |
| --- | --- | --- | --- |
| b. | 3 | d. | 5 |

ANS:	D

5 L/min  4 = 20. 20 + 20 = 0.40 FIO2.

PTS:	1	REF:	Page 98

96.	A patient is receiving supplemental oxygen via a 30% air-entrainment mask set at 8 L/min. The respiratory therapist must increase the supplemental oxygen to 40%. To maintain the same total output, the flow rate must be increased to _____ L/min.

| a. | 10 | c. | 16 |
| --- | --- | --- | --- |
| b. | 14 | d. | 18 |

ANS:	D

The output for the 30% air-entrainment mask set at 8 L/min is 72 L/min, because 30% has nine total parts multiplied by 8 L/min, which equals 72 L/min. The total output for 40% should, therefore, be 72 L/min. There are four total parts in 40% multiplied by an unknown flow that equals 72 L/min. 4X = 72. X is equal to 18 L/min.

PTS:	1	REF:	Page 98

97.	A patient is receiving supplemental oxygen at a flow of 4 L/min. Assessment reveals that there is clouding of the mask. What step should be taken to resolve this situation?

| a. | Change to a nasal cannula at 2 L/min. |
| --- | --- |
| b. | Increase the flow to 5 L/min. |
| c. | Change to a 50% air-entrainment mask. |
| d. | Increase the flow to 10% L/min. |

ANS:	B

The clouding of the mask shows there is CO2 retention in the mask. The minimum flow for a simple mask is 5 L/min. The other options would result in significant changes in the delivered FIO2.

REF Page 4-21

PTS:	1	REF:	Page 98

98.	A patient is being administered supplemental oxygen with a 40% air-entrainment mask being powered with a flow of 6 L/min. The manufacturer recommends a flow of 4 L/min. Which of the following will occur with this change?

1. There will be an increase in the delivered FIO2.

2. There will be an increase in the total flow.

3. There will be a decrease in the total flow.

4. There will be no change in the FIO2.

| a. | 1, 2 | c. | 3, 4 |
| --- | --- | --- | --- |
| b. | 2, 4 | d. | 1, 3 |

ANS:	B

The total flow from an air-entrainment mask will increase as the oxygen flow rate is increased. As long as the air:oxygen ratio is not changed, there will be no changes in the delivered FIO2.

REF Figure 4-21

PTS:	1	REF:	Page 99