Chapter 08: Assessment of Pulmonary Function

MULTIPLE CHOICE

1.	The accuracy of an instrument depends on which of the following?

1. The standard deviation of repeated measurements

2. The instrument’s linearity and frequency response

3. The instrument’s ability to reproduce a measurement

4. The instrument’s sensitivity to environmental conditions

| a. | 1 and 2 | c. | 2 and 3 |
| --- | --- | --- | --- |
| b. | 1 and 3 | d. | 2 and 4 |

ANS:	D

Accuracy can be explained as how closely a measured value is related to the correct value of the quantity measured. The accuracy of any instrument depends on its linearity and frequency response, its sensitivity to environmental conditions, and how well it is calibrated.

PTS:	1	REF:	Page 238

2.	The precision of an instrument depends on:

| a. | linearity. | c. | repeatability. |
| --- | --- | --- | --- |
| b. | calibration. | d. | frequency of response. |

ANS:	C

An instrument’s precision can be quantified by calculating the standard deviation of repeated measurements. Precision is the expression of an instrument’s ability to reproduce a measurement.

PTS:	1	REF:	Page 238

3.	Which of the following is a volume-collecting device?

| a. | Turbine flowmeter | c. | Dry rolling seal spirometer |
| --- | --- | --- | --- |
| b. | Thermal anemometer | d. | Pressure pneumotachograph |

ANS:	C

A typical example of a volume-collecting device is a dry rolling seal spirometer. The other devices are flow-sensing devices.

PTS:	1	REF:	Page 238

4.	Airflow is measured by which of the following?

| a. | “Hot-wire” anemometer | c. | Water-sealed spirometer |
| --- | --- | --- | --- |
| b. | Bellows spirometer | d. | Dry rolling seal spirometer |

ANS:	A

A typical example of a flow-sensing device is a hot-wire, or thermal, anemometer. The other devices are volume-collecting devices.

PTS:	1	REF:	Page 239

5.	The rotating motor-driven drum used to document the bell movement of a water-sealed spirometer is known as which of the following?

| a. | Spirograph | c. | Anemometer |
| --- | --- | --- | --- |
| b. | Kymograph | d. | Pneumotachograph |

ANS:	B

This is the definition of a kymograph.

PTS:	1	REF:	Page 239

6.	“An aluminum bell suspended by a chain-and-pulley mechanism with a weight that counterbalances the weight of the bell” describes which of the following?

| a. | Dry rolling spirometer | c. | Bellows spirometer |
| --- | --- | --- | --- |
| b. | Water-sealed spirometer | d. | Pneumotachograph |

ANS:	B

This is the description of a water-sealed spirometer.

PTS:	1	REF:	Page 239

7.	The visual picture of volume and flow measurements is known as which of the following?

| a. | Spirogram | c. | Kymograph |
| --- | --- | --- | --- |
| b. | Spirometer | d. | Turbine flowmeter |

ANS:	A

As a patient exhales into a water-sealed spirometer, the bell moves upward and the attached pen moves proportionately downward on graph paper, creating a spirogram.

PTS:	1	REF:	Page 240

8.	When a patient inhales while connected to a water-sealed spirometer, the bell moves _____ and the pen moves _____.

| a. | Up; up | c. | Down; up |
| --- | --- | --- | --- |
| b. | Up; down | d. | Down; down |

ANS:	C

Inhalation causes the bell to move downward and the pen to move upward.

PTS:	1	REF:	Page 240

9.	What speed should be used for the rotating drum of a water-sealed spirometer to measure volume changes during a forced vital capacity maneuver?

| a. | 16 mm/min | c. | 160 mm/min |
| --- | --- | --- | --- |
| b. | 32 mm/min | d. | 1920 mm/min |

ANS:	D

The fastest speed (1920 mm/min) is used for recording volume changes during forced vital capacity maneuvers.

PTS:	1	REF:	Page 240

10.	Speed settings of 32 and 160 mm/min on a water-sealed spirometer can be used to measure which of the following?

1. Minute ventilation

2. Peak expiratory flows

3. Maximum voluntary ventilation

4. Forced vital capacity maneuvers

| a. | 1 and 2 | c. | 2 and 4 |
| --- | --- | --- | --- |
| b. | 1 and 3 | d. | 3 and 4 |

ANS:	B

The slower speeds (32 and 160 mm/min) are used for measuring tidal volume, minute ventilation, and maximum voluntary ventilation.

PTS:	1	REF:	Page 239

11.	The bell factor for a 9-L water-sealed spirometer bell is _____ mL/mm.

| a. | 13.5 | c. | 20.93 |
| --- | --- | --- | --- |
| b. | 29.3 | d. | 41.73 |

ANS:	C

The bell factor is the number of milliliters of gas that must be displaced to cause the kymographic pen to move 1 mm. The bell factor for a 9-L bell is 20.93 mL/mm.

PTS:	1	REF:	Page 239

12.	What is the tidal volume measurement when the pen of a 9-L water-sealed spirometer moves 25 mm?

| a. | 0.338 L | c. | 0.843 L |
| --- | --- | --- | --- |
| b. | 0.523 L | d. | 1.043 L |

ANS:	B

The bell factor for a 9-L water-sealed spirometer is 20.93 mL/mm. Therefore, 20.93 mL/1 mm = X    mL/25 mm; X = 523 mL, or 0.523 L.

PTS:	1	REF:	Page 239

13.	What is the volume displacement when the pen of a 13.5-L water-sealed spirometer moves 28 mm?

| a. | 0.378 L | c. | 0.586 L |
| --- | --- | --- | --- |
| b. | 0.563 L | d. | 1.168 L |

ANS:	D

The bell factor for a 13.5-L water-sealed spirometer is 41.73 mL/mm. Therefore, 41.73 mL/1 mm = X    mL/28 mm; X = 1168 mL, or 1.168 L.

PTS:	1	REF:	Page 239

14.	Which of the following is true concerning the Stead-Wells spirometer?

| a. | Counterweights are necessary. |
| --- | --- |
| b. | The pen and bell move in opposite directions. |
| c. | The pen recorder is attached directly to the bell. |
| d. | On exhalation, the bell and pen move downward. |

ANS:	C

The pen recorder for the Stead-Wells spirometer is attached directly to the bell, so as the bell moves upward during exhalation, the pen also moves upward.

PTS:	1	REF:	Page 240

15.	________ is used to measure lung-volume changes by collecting exhaled gas into an expandable wedge-shaped device.

| a. | A bellows spirometer | c. | A vortex pneumotachometer |
| --- | --- | --- | --- |
| b. | A water-sealed spirometer | d. | A screen pneumotachometer |

ANS:	A

The expandable wedge-shaped device is a bellows, making this a bellows spirometer.

PTS:	1	REF:	Page 240

16.	Which statement is true concerning thermal flowmeters?

| a. | Gas flow can be measured bidirectionally. |
| --- | --- |
| b. | Gas velocity affects the amount of cooling. |
| c. | Cooling increases the resistance of a heated wire. |
| d. | Cooling decreases the resistance of a thermistor bead. |

ANS:	B

Gas flow is calculated by the amount of power needed to maintain the temperature of the heating element. A high gas velocity will cool the heating element, requiring more power to maintain the temperature of the thermistor.

PTS:	1	REF:	Page 241

17.	Which of the following devices is bidirectional?

| a. | Turbine flowmeter | c. | Vortex ultrasonic flowmeter |
| --- | --- | --- | --- |
| b. | Screen pneumotachograph | d. | Variable-orifice pneumotachograph |

ANS:	D

The variable-orifice pneumotachograph is a disposable, bidirectional, flow-measuring device that uses a flexible obstruction with a variable area for measuring flow as a function of the pressure differential generated by the obstruction. Turbine flowmeters, thermal anemometers, and vortex ultrasonic flowmeters are all unidirectional.

PTS:	1	REF:	Page 243

18.	Peak flowmeters are calibrated in which of the following units?

| a. | liters/second | c. | milliliters/second |
| --- | --- | --- | --- |
| b. | liters/minute | d. | millimeters/minute |

ANS:	B

Peak flowmeters operate by measuring the gas flow against a rotating vane. Peak flowmeters are usually calibrated in liters per minute.

PTS:	1	REF:	Page 241

19.	The principle that gas flow is proportional to the pressure drop across a known resistance is used by which of the following devices?

| a. | Turbine flowmeter | c. | Fleisch pneumotachograph |
| --- | --- | --- | --- |
| b. | Thermal flowmeter | d. | Ultrasonic pneumotachograph |

ANS:	C

All pneumotachographs, except the ultrasonic pneumotachograph, operate on the principle that gas flow through them is proportional to the pressure drop that occurs as the gas flows across a known resistance.

PTS:	1	REF:	Page 243

20.	Which type of flow-measuring device uses struts to create a partial obstruction to gas flow, causing whirlpools to be produced when gas flows through it?

| a. | Monel pneumotachograph | c. | Vortex ultrasonic flowmeter |
| --- | --- | --- | --- |
| b. | Ceramic pneumotachograph | d. | Variable-orifice pneumotachograph |

ANS:	C

The vortex ultrasonic flowmeter uses struts to create a partial obstruction to gas flow. As gases flow past these struts, whirlpools (or vortices) are produced. The frequency at which these whirlpools are produced is related to gas flow through the struts.

PTS:	1	REF:	Page 243

21.	Which type of pneumotachograph uses the Doppler effect to measure airflow velocity?

| a. | Fleisch | c. | Nonvortex |
| --- | --- | --- | --- |
| b. | Screen type | d. | Variable-orifice |

ANS:	C

The nonvortex ultrasonic pneumotachograph estimates airflow by projecting pulsed sound waves along the longitudinal axis of the flowmeter. The theory is that the speed of the ultrasonic wave transmission is influenced by the rate of gas flow through the device, which is the Doppler effect.

PTS:	1	REF:	Page 243

22.	A screen pneumotachograph is:

| a. | built with metal struts. | c. | bidirectional and very accurate. |
| --- | --- | --- | --- |
| b. | expensive and non-disposable. | d. | heated to prevent condensation. |

ANS:	D

The screen pneumotachograph is similar to Fleisch-type devices in that a heating element is incorporated to prevent water condensation on the metal screens.

PTS:	1	REF:	Page 243

23.	The standard for repeatability for forced expiratory volume in 1 second (FEV1) is ___ mL.

| a. | 100 | c. | 200 |
| --- | --- | --- | --- |
| b. | 150 | d. | 250 |

ANS:	B

The repeatability standard for FEV1 is 150 mL.

PTS:	1	REF:	Pages 243-245

24.	The best method for accurately measuring the thoracic gas volume of a patient with emphysema is:

| a. | helium dilution. | c. | pressure plethysmography. |
| --- | --- | --- | --- |
| b. | nitrogen washout. | d. | chemiluminescence monitoring. |

ANS:	C

Inert gas techniques, such as helium dilution and nitrogen washout, can measure gas volumes only in communicating airways. Body plethysmography measures all of the volume in the thorax. Therefore, inert gas measurements of functional residual capacity (FRC) for a patient with chronic obstructive pulmonary disease and severe air trapping will be lower than functional residual capacity measurements obtained during body plethysmography.

PTS:	1	REF:	Page 246

25.	Which of the following represents the relationship used by body plethysmography to calculate functional residual capacity?

| a. | VFRC = V ÷ P  (PB – PH2O) | c. | VFRC = V ÷ T  (PB – PH2O) |
| --- | --- | --- | --- |
| b. | VFRC = P ÷ V  (PB – PH2O) | d. | VFRC = (PB – PH2O)  V |

ANS:	A

VFRC = V ÷ P  (PB – PH2O), where V and P are the box volume and alveolar pressure changes measured during the panting maneuver, PB is the ambient barometric pressure, and PH2O is the water vapor pressure (assume 47 mm Hg for 37° C).

PTS:	1	REF:	Page 246

26.	The end point of a helium dilution test occurs when:

| a. | nitrogen is washed out. |
| --- | --- |
| b. | there is no more helium. |
| c. | the change in pressure remains steady. |
| d. | the helium percentage remains steady for 2 minutes. |

ANS:	D

The end point of the helium dilution test is reached when the helium percentage remains steady for 2 minutes, indicating that the helium is equilibrated between the spirometer and the patient’s lungs.

PTS:	1	REF:	Page 247

27.	One type of ventilatory pattern necessary for calculating functional residual capacity through the use of a body plethysmograph is:

| a. | deep breathing with a breath hold. | c. | deep, rapid breathing. |
| --- | --- | --- | --- |
| b. | rapid, shallow breathing. | d. | slow, deep breathing. |

ANS:	B

Mouth pressure and box pressure changes are measured during tidal breathing and panting maneuvers.

PTS:	1	REF:	Page 246

28.	Which of the following equipment is necessary to perform a nitrogen washout test?

| a. | Pneumotachometer | c. | Drying column |
| --- | --- | --- | --- |
| b. | Carbon dioxide (CO2) absorber | d. | Pressure transducer |

ANS:	A

See Figure 8-10, which shows the breathing circuit for performing a nitrogen washout test.

PTS:	1	REF:	Page 247, Figure 8-11

29.	Body plethysmography is based on which of the following gas laws?

| a. | Boyle’s | c. | Dalton’s |
| --- | --- | --- | --- |
| b. | Charles’ | d. | Gay-Lussac’s |

ANS:	A

Body plethysmography measures mouth pressure and box pressure changes during tidal volume and panting. The microprocessor unit calculates the functional residual capacity from empirically derived pressure-volume relationships. It is Boyle’s law that explains the relationship between gas volume and pressure when temperature is held constant.

PTS:	1	REF:	Page 247

30.	A comparison of functional residual capacity measured by body plethysmography with that measured by nitrogen washout reveals that the measurement obtained during the nitrogen washout test is less than that during body plethysmography. This discrepancy is most likely caused by which of the following?

| a. | Poor technique | c. | Restrictive lung disease |
| --- | --- | --- | --- |
| b. | Severe air-trapping | d. | Directional valve malfunction |

ANS:	B

Severe air-trapping will produce lower functional residual capacity results than body plethysmography because the nitrogen washout test measures gas volumes only in communicating airways.

PTS:	1	REF:	Page 247

31.	Aneroid manometers use ________ to measure pressure.

| a. | a vacuum chamber with a flexible diaphragm |
| --- | --- |
| b. | capacitance and changes in output voltage |
| c. | changes in inductance of two coils |
| d. | a Wheatstone bridge circuit |

ANS:	A

An aneroid manometer consists of a vacuum chamber with a flexible cover or diaphragm that flexes when pressure is applied to it. Electromechanical transducers use variable capacitance; strain gauge pressure transducers use Wheatstone bridge circuits; and variable inductance transducers use the magnetic flux between two coils to measure pressure.

PTS:	1	REF:	Page 250

32.	Which type of pressure-measurement device is best suited to measure arterial blood pressure through an arterial catheter?

| a. | Bourdon gauge | c. | Variable inductance |
| --- | --- | --- | --- |
| b. | Aneroid manometer | d. | Variable capacitance |

ANS:	C

Variable inductance and strain gauge pressure transducers are commonly used for measuring respiratory and cardiovascular pressures because they respond quickly to pressure changes.

PTS:	1	REF:	Page 250

33.	Which type of pressure-measurement device is used to measure a patient’s maximum inspiratory pressure?

| a. | Variable inductance | c. | Plethysmograph |
| --- | --- | --- | --- |
| b. | Aneroid manometer | d. | Strain gauge |

ANS:	B

Aneroid manometers are used as independent units for instantaneous pressure measurements, such as maximum inspiratory and expiratory pressures.

PTS:	1	REF:	Page 250

34.	Which of the following is true concerning the triple-lumen esophageal catheter used with the BICORE CP-100 Pulmonary Monitoring System?

| a. | It can measure intrapleural pressure indirectly. |
| --- | --- |
| b. | Its ideal position is the upper third of the esophagus. |
| c. | When the catheter enters the stomach, a negative pressure is recorded. |
| d. | The catheter can be positioned by observing the markings on its side. |

ANS:	A

Esophageal pressure is used clinically to estimate intrapleural pressure.

PTS:	1	REF:	Page 250

35.	The bedside respiratory mechanics of mechanically ventilated patients can be measured with which of the following?

1. VenTrak system

2. Tissot spirometer

3. BICORE CP-100 system

4. Stead-Wells spirometer

| a. | 1 and 2 | c. | 2 and 3 |
| --- | --- | --- | --- |
| b. | 1 and 3 | d. | 2 and 4 |

ANS:	B

Stand-alone devices for the measurement of bedside respiratory mechanics include the BICORE CP-100 and the VenTrak system.

PTS:	1	REF:	Page 251

36.	“Opposition to the flow of an alternating current” is the definition of:

| a. | electrical resistance. | c. | electrical impedance. |
| --- | --- | --- | --- |
| b. | variable capacitance. | d. | variable inductance. |

ANS:	C

Electrical impedance, which is opposition to the flow of an alternating current, is determined by the resistance and capacitance of the circuit through which the current must pass.

PTS:	1	REF:	Pages 251-252

37.	Impedance plethysmographs operate by using _____-frequency, _____-amplitude electrical current.

| a. | Low; low | c. | Low; high |
| --- | --- | --- | --- |
| b. | High; low | d. | High; high |

ANS:	B

A constant high-frequency, low-amplitude electrical current is passed between the two electrodes, and the return voltage is used to calculate the impedance.

PTS:	1	REF:	Page 252

38.	Which of the following might cause an apnea monitor to fail to alert the clinician to the presence of apnea?

| a. | Tachycardia | c. | Upper airway obstruction |
| --- | --- | --- | --- |
| b. | Changes in blood volume | d. | Inadequate Sensitivity settings |

ANS:	C

Recent evidence suggests that bradycardia and upper airway obstruction might cause apnea monitors to fail to signal apnea.

PTS:	1	REF:	Page 252

39.	The parents of a newborn sent home with an apnea monitor call the respiratory therapist to report several periods of apnea as noted by the monitor’s alarm system. The respiratory therapist should check which of the following to ensure that the baby’s movement was not the cause of these alarms?

| a. | Sensitivity of the monitor | c. | Amount of electrical current |
| --- | --- | --- | --- |
| b. | Placement of the electrodes | d. | Frequency of the vibrations |

ANS:	A

The sensitivity of the impedance pneumograph can be adjusted by the operator to prevent false alarms because of changes in impedance caused by movement instead of by changes in respiration.

PTS:	1	REF:	Page 252

40.	Which of the following are factors that influence the partial pressure of alveolar carbon dioxide (PACO2)?

1. Diffusion

2. Ventilation

3. CO2 production

4. Perfusion of the lungs

| a. | 1 and 2 | c. | 1, 3, and 4 |
| --- | --- | --- | --- |
| b. | 2 and 3 | d. | 1, 2, 3, and 4 |

ANS:	D

Four major factors influence the PACO2—CO2 production, perfusion of the lungs, diffusion, and ventilation. The balance of these four components represents the total transport and elimination of CO2.

PTS:	1	REF:	Page 266

41.	While working in the neonatal intensive care unit, a respiratory therapist notices that a baby with an apnea monitor is not breathing, but the apnea alarm is not ringing. Which of the following might cause this failure?

1. Tachypnea

2. Bradycardia

3. Excessive movement

4. Upper airway obstruction

| a. | 1 and 2 | c. | 1 and 3 |
| --- | --- | --- | --- |
| b. | 2 and 4 | d. | 3 and 4 |

ANS:	B

Bradycardia and upper airway obstruction can cause apnea monitors to fail to signal apnea, because with bradycardia, cardiac oscillations and intrathoracic blood-volume changes cause an increase in impedance, which is sensed as part of the respiratory cycle. Continued respiratory efforts in the presence of upper airway obstruction are sensed as normal respiratory efforts.

PTS:	1	REF:	Page 252

42.	What type of device is used in an ICU to detect respiratory muscle fatigue?

| a. | Aneroid manometer | c. | Inductive plethysmograph |
| --- | --- | --- | --- |
| b. | Pressure plethysmograph | d. | Electromechanical transducer |

ANS:	C

Respiratory inductive plethysmography has been used clinically as a means of monitoring the breathing patterns of patients in ICUs. In the ICU, it has been used primarily to identify uncoordinated thoracoabdominal movements that are associated with respiratory muscle fatigue or failure.

PTS:	1	REF:	Page 277

43.	The principle that the respiratory system moves with 2 degrees of freedom describes the basis for which of the following?

| a. | Inductive plethysmography | c. | Strain gauge transducer |
| --- | --- | --- | --- |
| b. | Volume plethysmography | d. | Nitrogen washout |

ANS:	A

Inductive plethysmography is based on the principle that the respiratory system moves with 2 degrees of freedom. That is, it consists of two moving parts: the rib cage and the abdomen. During inspiration, the rib cage moves outward as the lungs expand in the thorax. Simultaneously, the abdomen is displaced outward by the downward movement of the diaphragm.

PTS:	1	REF:	Page 252

44.	As measured by inductive plethysmography, rib cage movement plus abdominal movement equals:

| a. | tidal volume. | c. | expiratory reserve volume. |
| --- | --- | --- | --- |
| b. | inspiratory capacity. | d. | functional residual capacity. |

ANS:	A

Because the two moving parts in the respiratory system, the rib cage and the abdomen, are arranged in a series, the sum of the two displacements can be used to calculate the air volume inspired.

PTS:	1	REF:	Page 253

45.	Which device would be appropriate for monitoring respiratory rate and tidal volume during a sleep study?

| a. | Capnograph | c. | Pneumotachograph |
| --- | --- | --- | --- |
| b. | Mass spectrometer | d. | Inductive plethysmograph |

ANS:	D

Respiratory inductive plethysmography has been used extensively in research on respiratory muscle function. It has been used as a means of monitoring breathing patterns of patients in sleep laboratories.

PTS:	1	REF:	Page 253

46.	Which of the following values is not measured directly?

| a. | FEV1 | c. | Peak expiratory flows |
| --- | --- | --- | --- |
| b. | Airway resistance (Raw) | d. | Maximum inspiratory pressure |

ANS:	B

FEV1 and peak expiratory flows are routine measurements during spirometry. Maximum inspiratory pressure is a common airway pressure measurement made on spontaneously breathing patients. Raw is calculated by subtracting plateau pressure (Pplateau) from peak inspiratory pressure and dividing the resultant pressure by the airflow.

PTS:	1	REF:	Page 257

47.	Which of the following variables can be measured directly?

| a. | Lung compliance | c. | Raw |
| --- | --- | --- | --- |
| b. | Work of breathing | d. | Pplateau |

ANS:	D

Pplateau is measured during mechanical ventilation and represents the amount of pressure needed to maintain the tidal volume within the patient’s lungs during a period of no gas flow.

PTS:	1	REF:	Page 256

48.	To set up a pulmonary function laboratory that can measure all standard and dynamic lung volumes, which of the following equipment is necessary?

| a. | Screen pneumotachograph and inductive plethysmograph |
| --- | --- |
| b. | Fleisch pneumotachograph and peak flowmeter |
| c. | Vortex ultrasonic flowmeter and inductive plethysmograph |
| d. | Stead-Wells spirometer and constant-volume plethysmograph |

ANS:	D

The Stead-Wells spirometer measures volumes and flows, whereas the constant-volume plethysmograph (the most common plethysmograph used) measures functional residual capacity. These two devices are able to measure all the volumes and flows necessary for full pulmonary function studies. The screen, Fleisch, and vortex ultrasonic pneumotachographs are unidirectional and are not able to measure inspiratory and expiratory flow simultaneously. Inductive plethysmography will provide respiratory rate and volume only, and is used in sleep laboratories.

PTS:	1	REF:	Pages 243-245

49.	The most appropriate lung volume to measure for a mechanically ventilated patient is _____ volume.

| a. | minute | c. | expiratory reserve |
| --- | --- | --- | --- |
| b. | residual | d. | inspiratory reserve |

ANS:	A

At the bedside, the most commonly measured lung volume is the expired minute volume.

PTS:	1	REF:	Page 254

50.	Newer ventilators can make tidal volume and minute volume measurements by incorporating which of the following?

| a. | Capnograph | c. | Flow transducer |
| --- | --- | --- | --- |
| b. | Anemometer | d. | Pressure transducer |

ANS:	C

In most new ventilators, a flow transducer is incorporated into the system design to give continual updates on the tidal volume and minute volume.

PTS:	1	REF:	Page 254

51.	If a patient’s minute volume is 8.5 L/min and his respiratory rate is 15 breaths/min, his average tidal volume is ____ L.

| a. | 0.125 | c. | 0.060 |
| --- | --- | --- | --- |
| b. | 0.570 | d. | 5.700 |

ANS:	B

Minute volume = tidal volume  respiratory rate.

Tidal volume = minute volume ÷ respiratory rate.

Tidal volume = 8.5 L/min ÷ 15 breaths/min.

Tidal volume = 0.570 L.

PTS:	1	REF:	Page 253

52.	What is the minute volume of a patient if her respiratory rate is 12 breaths/min and her tidal volume is 750 mL?

| a. | 0.06 L | c. | 6.25 L |
| --- | --- | --- | --- |
| b. | 9.00 L | d. | 9.00 L |

ANS:	D

Minute volume = tidal volume  respiratory rate.

Minute volume = 0.75 L 12 breaths/min.

Minute volume = 9 L.

PTS:	1	REF:	Page 253

53.	Increased alveolar ventilation is associated with which of the following?

| a. | Hyperthermia | c. | Decreased metabolic rate |
| --- | --- | --- | --- |
| b. | Bronchospasm | d. | Respiratory muscle fatigue |

ANS:	A

Hyperthermia increases the metabolic demands of the tissues and increases alveolar ventilation.

PTS:	1	REF:	Page 254

54.	Which of the following devices can a patient with asthma use in the home to monitor daily variations in Raw?

| a. | Thermal flowmeter | c. | Screen pneumotachograph |
| --- | --- | --- | --- |
| b. | Fleisch pneumotachograph | d. | Vortex ultrasonic flowmeter |

ANS:	C

Screen pneumotachographs are inexpensive, disposable, operate at ambient temperature, and do not require a heating element. The Fleisch pneumotachograph is affected by turbulent airflow, making accurate peak flow readings difficult.

PTS:	1	REF:	Page 243

55.	Secretions in the airway of a mechanically ventilated patient could be indicated by a(n):

| a. | increase in inspiratory capacity. |
| --- | --- |
| b. | increase in breathing frequency. |
| c. | increase in peak inspiratory pressure. |
| d. | decrease in forced expiratory flow rates. |

ANS:	C

An increase in airway resistance, as with secretions in the airway, will cause an elevation in the peak inspiratory pressure.

PTS:	1	REF:	Page 257

56.	Vital capacity is the sum of which of the following?

1. Residual volume

2. Inspiratory capacity

3. Expiratory reserve volume

4. Inspiratory reserve volume

5. Functional residual capacity

| a. | 2 and 3 | c. | 2 and 5 |
| --- | --- | --- | --- |
| b. | 1 and 4 | d. | 3 and 4 |

ANS:	A

The vital capacity is equal to the sum of the inspiratory capacity plus the expiratory residual volume.

PTS:	1	REF:	Page 253, Figure 8-18

57.	Expiratory residual volume plus residual volume equals:

| a. | functional residual capacity. | c. | total lung capacity. |
| --- | --- | --- | --- |
| b. | inspiratory capacity. | d. | vital capacity. |

ANS:	A

Expiratory residual volume plus residual volume equals functional residual capacity.

PTS:	1	REF:	Page 253, Figure 8-18

58.	Total lung capacity is a result of:

| a. | vital capacity plus residual volume. |
| --- | --- |
| b. | inspiratory capacity plus residual volume. |
| c. | inspiratory capacity plus expiratory residual volume. |
| d. | inspiratory reserve volume plus functional residual capacity. |

ANS:	A

The total lung capacity is the sum of the vital capacity and the residual volume.

PTS:	1	REF:	Page 253, Figure 8-18

59.	Patients with restrictive pulmonary disease demonstrate a decrease in which of the following measurements?

1. Vital capacity

2. Total lung capacity

3. Forced expiratory flow from 25% to 75% of the vital capacity (FEF25-75)

4. Ratio of forced expiratory volume in 1 second to vital capacity (FEV1/VC)

| a. | 1 and 2 | c. | 1 and 4 |
| --- | --- | --- | --- |
| b. | 2 and 3 | d. | 3 and 4 |

ANS:	A

Patients with restrictive pulmonary disease can have decreases in vital capacity and total lung capacity.

PTS:	1	REF:	Page 254

60.	The results of pulmonary function testing show a normal total lung capacity and decreased, but reversible, FEV1 and FEV25-75. The patient is a 22-year-old, nonsmoking man with a history of intermittent “noisy breathing.” Which of the following is most likely the cause of his problem?

| a. | Asthma | c. | Pulmonary fibrosis |
| --- | --- | --- | --- |
| b. | Emphysema | d. | Severe restrictive pulmonary disease |

ANS:	A

Asthma is a disease that shows reversibility in the FEV1 and FEV25-75. This reversibility, coupled with the fact that he is an otherwise healthy, nonsmoking, 22- year-old with a history of intermittent noisy breathing, makes asthma the best choice as the most likely cause of the problem.

PTS:	1	REF:	Page 257

61.	If the chest wall compliance of a mechanically ventilated patient decreases, which of the following will occur?

| a. | Maximum inspiratory pressure will decrease. |
| --- | --- |
| b. | Peak inspiratory pressure will decrease. |
| c. | Pplateau will increase. |
| d. | Tidal volume will increase. |

ANS:	C

When chest wall compliance decreases, there will be an increase in maximum inspiratory pressure and Pplateau.

PTS:	1	REF:	Page 257

62.	Which device can be incorporated into a ventilator to provide continuous recordings of airway pressure during the breathing cycle?

| a. | Variable capacitance transducer | c. | Strain gauge transducer |
| --- | --- | --- | --- |
| b. | Inductive plethysmograph | d. | Aneroid manometer |

ANS:	C

Peak inspiratory pressure can be derived from continuous recordings of airway pressure that can be made during the breathing cycle by using a strain gauge transducer.

PTS:	1	REF:	Page 256

63.	The amount of pressure required by a ventilator to maintain tidal volume when there is no gas flow is known as:

| a. | Pplateau. | c. | peak expiratory pressure. |
| --- | --- | --- | --- |
| b. | peak inspiratory pressure. | d. | Maximum inspiratory pressure. |

ANS:	A

Pplateau represents the amount of pressure needed to maintain the tidal volume within the patient’s lungs during static conditions and is the total pressure required to overcome only elastic forces.

PTS:	1	REF:	Page 256

64.	Which device is appropriate for assessing a patient’s response to bronchodilator therapy at the bedside?

| a. | Stead-Wells spirometer | c. | Inductive plethysmograph |
| --- | --- | --- | --- |
| b. | Mainstream capnograph | d. | Rotating-vane flowmeter |

ANS:	D

Rotating-vane flowmeters are used as handheld peak flowmeters and at the bedside to assess patient response to bronchodilator therapy.

PTS:	1	REF:	Page 241

65.	Which of the following is a complication of performing peak expiratory flows maneuvers?

| a. | Increased static pressure | c. | Decreased chest compliance |
| --- | --- | --- | --- |
| b. | Increased airway resistance | d. | Decreased mean airway pressure |

ANS:	B

Increased airway resistance, due to bronchoconstriction or airway collapse, is a complication of forced exhalation required to perform peak expiratory flows maneuvers.

PTS:	1	REF:	Page 256, Clinical Practice Guideline 8-1

66.	Occlusion of the expiratory valve of the ventilator at the end of a tidal inspiration enables measurement of which of the following?

| a. | Pplateau | c. | Maximum inspiratory pressure |
| --- | --- | --- | --- |
| b. | Peak airway pressure | d. | Maximum expiratory pressure |

ANS:	A

Pplateau is measured during mechanical ventilation by temporarily occluding the expiratory valve of the ventilator at the end of a tidal inspiration and noting the new pressure level.

PTS:	1	REF:	Page 256

67.	What is the pressure reading on the anemometer of a ventilator during an inspiratory hold maneuver?

| a. | Pplateau. | c. | Mean inspiratory pressure |
| --- | --- | --- | --- |
| b. | Peak inspiratory pressure | d. | Maximum inspiratory pressure |

ANS:	A

The inspiratory hold maneuver shows the pressure to maintain the tidal volume within the patient’s lungs during a period of no gas flow; this is the Pplateau.

PTS:	1	REF:	Page 256

68.	Which gas law is applicable when considering Raw?

| a. | Boyle’s | c. | Poiseuille’s |
| --- | --- | --- | --- |
| b. | Charles’ | d. | Avogadro’s |

ANS:	C

Poiseuille’s law states that a twofold decrease in airway diameter results in a 16-fold increase in Raw. Boyle’s law and Charles’ law address the relationships among pressure, volume, and temperature. Avogadro’s law is concerned with the number of molecules in a volume of gas.

PTS:	1	REF:	Page 257

69.	Normal Raw of the respiratory system is _____ cm H2O/L/s.

| a. | 1 to 3 | c. | 4 to 8 |
| --- | --- | --- | --- |
| b. | 2 to 5 | d. | 6 to 10 |

ANS:	B

Raw averages 2 to 5 cm H2O/L/s.

PTS:	1	REF:	Page 257

70.	Which of the following variables in the formula describing Poiseuille’s law can be altered in the clinical setting to improve Raw?

1. Radius of the airways

2. Viscosity of the inspired gas

3. Flow rate of the inspired gas

4. Temperature of the inspired gas

5. Length of the tracheobronchial tree

| a. | 1 and 3 | c. | 2, 3, and 5 |
| --- | --- | --- | --- |
| b. | 1, 4, and 5 | d. | 1, 2, and 3 |

ANS:	D

Poiseuille’s law is represented by the formula Raw = [8  gas flow  gas viscosity  airway length] ÷ [ radius4]. Bronchospasm causes a decrease in the radius of the airways. Bronchodilators can reverse that process, thereby decreasing the Raw. High flow rates cause more turbulence and increased Raw. During mechanical ventilation, the flow rate of the gas being delivered can be changed by the respiratory therapist. Helium-oxygen (heliox) is used in respiratory care, not because of its viscosity, but because of its low density. An 80:20 mix of heliox is the least-dense gas available for human use, but it has a viscosity higher than room air. So, heliox does not improve Raw. It does, however, make its way around any airway obstruction.

PTS:	1	REF:	Page 257

71.	Which of the following pathophysiologic conditions can increase both the peak inspiratory pressure and the Pplateau of a ventilated patient?

| a. | Asthma | c. | Kyphoscoliosis |
| --- | --- | --- | --- |
| b. | Atelectasis | d. | Pulmonary interstitial fibrosis |

ANS:	A

An asthma exacerbation causes bronchospasm, which increases Raw, thereby triggering an increase in peak inspiratory pressure without an increase in Pplateau.

PTS:	1	REF:	Page 256, Clinical Practice Guideline 8-1

72.	The respiratory system compliance of a patient receiving mechanical ventilation when the returned tidal volume is 650 mL, the static pressure is 35 cm H2O, and the positive end-expiratory pressure is 8 cm H2O is _____ mL/cm H2O.

| a. | 0.04 | c. | 53.9 |
| --- | --- | --- | --- |
| b. | 18.6 | d. | 24 |

ANS:	D

Respiratory system compliance can be determined by dividing the tidal volume by the Pplateau – baseline Cstat = 650 ÷ (35 – 8) = 650 ÷ 27 = 24 mL/cm H2O.

PTS:	1	REF:	Page 258, Table 8-4

73.	A(n) ______ analyzer is an oxygen analyzer that uses paramagnetism to measure oxygen concentration.

| a. | Galvanic | c. | Paramagnetic |
| --- | --- | --- | --- |
| b. | Electrical | d. | Polarographic |

ANS:	C

The paramagnetic analyzer uses the physical principle of paramagnetism to measure oxygen concentration. Galvanic and polarographic analyzers use chemical reactions to measure oxygen concentration, and electrical analyzers use thermal conductivity.

PTS:	1	REF:	Pages 257-258

74.	A(n) ______ analyzer is an oxygen analyzer that can display both partial pressure (in millimeters of mercury) and the percentage of oxygen concentration.

| a. | Paramagnetic | c. | Electrical |
| --- | --- | --- | --- |
| b. | Galvanic | d. | Polarographic |

ANS:	A

The scale on a paramagnetic oxygen analyzer is calibrated to display both partial pressures (in millimeters of mercury) and the percent of oxygen concentration.

PTS:	1	REF:	Page 257

75.	Which type of oxygen analyzer should not be used in the presence of anesthetic gases?

| a. | Polarographic | c. | Electrical |
| --- | --- | --- | --- |
| b. | Paramagnetic | d. | Galvanic |

ANS:	C

An electrical analyzer uses a Wheatstone bridge setup and monitors the changes in current flow as it relates to oxygen concentrations. The Wheatstone bridge can generate a significant amount of heat and therefore is dangerous to use in the presence of flammable gases, such as anesthetic gases.

PTS:	1	REF:	Page 259

76.	Which type of oxygen analyzer will give erroneous concentration levels in the presence of contaminant gases?

| a. | Polarographic | c. | Galvanic |
| --- | --- | --- | --- |
| b. | Electrical | d. | Physical |

ANS:	B

Electrical analyzers use changes in current flow through wires to measure the concentration of oxygen. Contaminant gases can dissipate heat at different rates than oxygen and nitrogen, so, if gases other than oxygen and nitrogen are present, this can lead to erroneous fractional inspired oxygen (FIO2) levels.

PTS:	1	REF:	Page 259

77.	The oxygen analyzer that is no longer commonly used is the ____ analyzer.

| a. | Paramagnetic | c. | Electrical |
| --- | --- | --- | --- |
| b. | Galvanic | d. | Polarographic |

ANS:	A

The paramagnetic oxygen analyzer, such as a Beckman D-2, is no longer commonly used.

PTS:	1	REF:	Page 257

78.	Which oxygen analyzer uses a semipermeable polytetrafluoroethylene membrane to separate oxygen from a potassium hydroxide bath containing a gold cathode and a lead anode?

| a. | Paramagnetic | c. | Electrical |
| --- | --- | --- | --- |
| b. | Galvanic | d. | Polarographic |

ANS:	B

Within the galvanic cell there is a semipermeable membrane that separates the gas sample from a hydroxide solution with a gold cathode and a lead anode. The polarographic cell uses a platinum cathode and a silver anode immersed in a potassium hydroxide bath.

PTS:	1	REF:	Page 258

79.	Electrical analyzers operate by using which of the following?

| a. | Clark electrodes | c. | Thermal conductivity |
| --- | --- | --- | --- |
| b. | Pauling’s principle | d. | Oxygen-mediated chemical reaction |

ANS:	C

Electrical analyzers operate on the principle of thermal conductivity and use an electronic device called a Wheatstone bridge. The Clark electrode is used by the polarographic analyzer. Pauling’s principle is used in paramagnetic analyzers, and electrical analyzers use thermal conductivity to analyze oxygen concentration.

PTS:	1	REF:	Page 259

80.	Which of the following happen to the wire within a Wheatstone bridge when it is exposed to oxygen concentrations higher than room air?

1. Its resistance decreases.

2. Its resistance increases.

3. It cools down.

4. It heats up.

| a. | 1 and 4 | c. | 1 and 4 |
| --- | --- | --- | --- |
| b. | 2 and 3 | d. | 1 and 3 |

ANS:	D

The Wheatstone bridge contains two parallel wires that receive current flow from an external power source, usually a battery. One of these wires, which serves as the reference, is exposed to room air. The other wire is located in the sample chamber and is exposed to the gas being analyzed. If the sample gas contains a higher oxygen-to-nitrogen ratio than is present in room air, the sample wire cools and its resistance decreases, because oxygen is a better conductor of heat than nitrogen.

PTS:	1	REF:	Page 259

81.	Which analyzer would be most appropriate to use during an airlift transport of a patient?

| a. | Galvanic | c. | Clark electrode |
| --- | --- | --- | --- |
| b. | Electrical | d. | Polarographic |

ANS:	B

Electrical oxygen analyzers are unaffected by changes in pressure that occur with altitude changes.

PTS:	1	REF:	Page 259

82.	Which of the following types of oxygen analyzer can be placed in a ventilator circuit to continuously monitor oxygen concentration?

1. Polarographic

2. Paramagnetic

3. Galvanic

4. Physical

| a. | 1 and 4 | c. | 2 and 3 |
| --- | --- | --- | --- |
| b. | 1 and 3 | d. | 2 and 4 |

ANS:	B

Galvanic and polarographic analyzers can be used for continuous monitoring of oxygen concentrations during mechanical ventilation.

PTS:	1	REF:	Page 258

83.	Nitric oxide is effective in the treatment of pulmonary hypertension in neonates because it causes which of the following?

| a. | Peripheral vasoconstriction | c. | Platelet adhesion |
| --- | --- | --- | --- |
| b. | Macrophage cytotoxicity | d. | Vasodilatation |

ANS:	D

Nitric oxide is used for the treatment of pulmonary hypertension in neonates because it promotes pulmonary vasodilatation. Other effects of nitric oxide include macrophage cytotoxicity and platelet adhesion.

PTS:	1	REF:	Page 260

84.	Two gases that can be falsely identified as nitrogen dioxide (NO2) are:

| a. | NH2 and HNO3 | c. | NO3 and HNO3 |
| --- | --- | --- | --- |
| b. | NH3 and HNO3 | d. | NH2 and nitrous oxide (N2O) |

ANS:	D

NH2 and N2O are two gases that can be falsely identified as NO2.

PTS:	1	REF:	Page 260

85.	A device that can measure NO2 uses which of the following?

| a. | A Wheatstone bridge | c. | Potassium hydroxide |
| --- | --- | --- | --- |
| b. | A catalytic converter | d. | An infrared (IR) spectroscope |

ANS:	B

NO2 levels can be measured indirectly by converting NO2 to nitric oxide (NO) with a catalytic or chemical converter, and then measuring the NO concentration.

PTS:	1	REF:	Page 260

86.	Which reactions take place near the sensing electrode of an electrochemical NO analyzer?

| a. | 2H2O  4H+ + 4e– + O2 | c. | 4NO + 2H2O + 2O2  4HNO3 |
| --- | --- | --- | --- |
| b. | O2 + 4H+ + 4e–  2H2O | d. | NO + 2H2O  HNO3 + 3H+ + 3e– |

ANS:	D

An electrochemical NO analyzer consists of three electrodes immersed in an electrolyte solution that contains a highly conductive concentrated acid or alkali solution. The electrodes are separated from the gas sample to be analyzed by a semipermeable membrane. NO and NO2, from the unknown gas sample, diffuse across a semipermeable membrane and react with the electrolyte solution near the sensing electrode. This generates electrons in the oxidation reaction: NO + 2H2O  HNO3 + 3H+ + 3e–.

PTS:	1	REF:	Page 260

87.	Which of the following gases affect the accuracy of NO/NO2 electrochemical monitors?

1. NH2

2. N2O

3. CO2

4. Carbon monoxide (CO)

5. Nitrogen (N2)

| a. | 1, 4, and 5 | c. | 2, 3, and 5 |
| --- | --- | --- | --- |
| b. | 1, 3, and 4 | d. | 3 and 4 |

ANS:	B

The accuracy of NO/NO2 electrochemical monitors can be altered by the presence of background gases—such as CO2, CO, and NH2—that have lower oxidation potentials than the sensor potential.

PTS:	1	REF:	Pages 260-261

88.	Which type of CO2 monitor is most appropriate for a critical care setting?

| a. | IR spectroscopy | c. | Mass spectroscopy |
| --- | --- | --- | --- |
| b. | Raman spectroscopy | d. | Clark electrode |

ANS:	A

IR spectroscopy is currently the method of choice in critical care settings, and mass spectroscopy is most often used in surgical suites.

PTS:	1	REF:	Page 262

89.	Which of the following are associated with sidestream capnographs?

1. Delayed reporting times

2. Addition of dead space to the airway

3. Increased risk of inappropriate extubation

4. Increased risk of plugging of the sample tubing

| a. | 1 and 2 | c. | 2 and 4 |
| --- | --- | --- | --- |
| b. | 1 and 3 | d. | 3 and 4 |

ANS:	B

Sidestream capnographs show a slight delay between sampling and reporting times because of the time required to transport the sample from the airway to the sample chamber. The plastic tube that transports sample gas from the airway to the analyzer is prone to plugging by water and secretions, and therefore can lead to erroneous readings.

PTS:	1	REF:	Page 264

90.	The gases that have the same molecular weight and cannot be separated on the basis of weight alone during mass spectroscopy include which of the following?

| a. | N2 and O2 | c. | CO2 and N2 |
| --- | --- | --- | --- |
| b. | N2O and O2 | d. | N2O and CO2 |

ANS:	D

CO2 and N2O have the same molecular weight and are not able to be separated on the basis of molecular weight, which can lead to erroneous readings. This problem is overcome by ionizing N2O to N2O+ and CO2 to C+.

PTS:	1	REF:	Page 265

91.	A normal capnogram contains the most CO2 during which of the following phases?

| a. | 1 | c. | 3 |
| --- | --- | --- | --- |
| b. | 2 | d. | 4 |

ANS:	C

During phase 1, there is essentially no CO2 because the gas is from the patient’s dead space. In phase 2, CO2 from the alveoli mixes with gas from the smaller airways, causing the CO2 concentration to rise. During phase 3, the CO2 concentration curve remains at its highest level as alveolar gas is exhaled. Inspiration causes the measured exhaled CO2 level to fall to 0; this is phase 4.

PTS:	1	REF:	Page 266

92.	End-tidal CO2 is measured with a capnograph during the _____ of phase _____.

| a. | end; 3 | c. | middle; 2 |
| --- | --- | --- | --- |
| b. | end; 4 | d. | beginning; 3 |

ANS:	A

The end of phase 3 has the highest amount of exhaled CO2.

PTS:	1	REF:	Page 266

93.	What is the normal relationship between the partial pressure of arterial carbon dioxide (PaCO2) and partial pressure of end-tidal CO2 (PETCO2)?

| a. | PaCO2 and PETCO2 are equal. |
| --- | --- |
| b. | PaCO2 is 10 mm Hg more than PETCO2. |
| c. | PaCO2 is 4 to 6 mm Hg less than PETCO2. |
| d. | PaCO2 is 4 to 6 mm Hg more than PETCO2. |

ANS:	D

Although the partial pressure of end-tidal CO2 (PETCO2) should equal the partial pressure of arterial carbon dioxide (PaCO2), the PETCO2 is actually 4 to 6 mm Hg lower than the PaCO2.

PTS:	1	REF:	Page 266

94.	At the alveolar level, when ventilation occurs with reduced or no perfusion, this condition is known as which of the following?

| a. | Alveolar shunt | c. | Anatomic dead space |
| --- | --- | --- | --- |
| b. | Anatomic shunt | d. | Functional dead space |

ANS:	D

Ventilation with little or no perfusion at the level of the alveoli is known as functional dead space. An alveolar shunt occurs when alveoli are perfused but not ventilated. An anatomic shunt is the percentage of cardiac output that does not participate in gas exchange in the lung because of normal anatomy: Unoxygenated blood from the bronchopulmonary vein (lungs) and the thebesian veins (heart) empty directly into the left heart. Anatomic dead space includes all the conducting airways where no gas exchange occurs.

PTS:	1	REF:	Page 266

95.	A 24-year-old skier with a broken femur has been in the postanesthesia care unit for 24 hours and is 12 hours post-extubation with no supplemental oxygen. He is awake, alert, and oriented, and he reports shortness of breath. A capnograph shows decreased PETCO2. This patient’s current status could be caused by which of the following?

| a. | Respiratory center depression | c. | Muscular paralysis |
| --- | --- | --- | --- |
| b. | Pulmonary embolism | d. | Alveolar shunt |

ANS:	B

Long-bone fractures put a patient at risk for the development of pulmonary emboli. The decreased PETCO2 is showing that physiologic dead space has increased because of the blockage of pulmonary arterioles.

PTS:	1	REF:	Page 276

96.	Which of the following PETCO2 readings is consistent with a drug overdose?

| a. | 0.3% | c. | 5.3% |
| --- | --- | --- | --- |
| b. | 2.5% | d. | 10.0% |

ANS:	D

Normal expired air contains approximately 4% to 6% CO2. A drug overdose depresses the respiratory center, leading to a higher than normal PETCO2.

PTS:	1	REF:	Page 265

97.	Capnography is widely used to determine:

| a. | hemodynamic status. | c. | placement of an endotracheal tube. |
| --- | --- | --- | --- |
| b. | adequacy of cerebral blood flow. | d. | response to bronchodilator therapy. |

ANS:	C

Capnography is an excellent method to determine tracheal, rather than esophageal, intubation because only the lungs produce CO2.

PTS:	1	REF:	Page 268, Clinical Practice Guideline 8-2

98.	Measuring the energy needs of patients is known as which of the following?

| a. | Oximetry | c. | Capnography |
| --- | --- | --- | --- |
| b. | Calorimetry | d. | Spectroscopy |

ANS:	B

Calorimetry is a method of determining energy expenditure by using the measurements of the amount of heat radiated and absorbed by an organism.

PTS:	1	REF:	Page 272

99.	Which type of calorimeter is best suited to measure the energy expenditure of a mechanically ventilated patient with a heated humidifier?

| a. | IR | c. | Closed-circuit |
| --- | --- | --- | --- |
| b. | Open-circuit | d. | Double-circuit |

ANS:	C

Closed-circuit calorimeters can be used for both spontaneously breathing and mechanically ventilated patients. Open-circuit devices can also be used for both spontaneously breathing and mechanically ventilated patients. However, special problems can arise when systems are used with mechanically ventilated patients. Some of these problems include fluctuations in FIO2, separation of the patient’s inspired and expired gases from the continuous gas flow provided by the ventilator, and handling of water vapor.

PTS:	1	REF:	Page 272

100.	Most calorimeters are unreliable when the oxygen concentration is above 50% to 60% because of which of the following?

| a. | Boyle’s law | c. | The Hamburger effect |
| --- | --- | --- | --- |
| b. | Charles’ law | d. | The Haldane transformation |

ANS:	D

Beyond FIO2s of 0.50 to 0.60, most systems are unreliable because of the Haldane transformation; therefore, any data derived from their use should accordingly be viewed skeptically.

PTS:	1	REF:	Page 275

101.	Which of the following factors will cause inaccurate calorimetry measurements on a mechanically ventilated patient?

| a. | FIO2 >0.5 | c. | Room temperature = 23° C |
| --- | --- | --- | --- |
| b. | Varying pressure | d. | Heat and moisture exchanger |

ANS:	A

Clinical studies have demonstrated that currently available systems cannot provide accurate and reproducible oxygenation consumption per minute (O2) measurements for patients breathing FIO2s greater than 0.5.

PTS:	1	REF:	Page 275

102.	Proteins contribute significantly to energy expenditure during which of the following states?

| a. | Severe sepsis | c. | Prolonged starvation |
| --- | --- | --- | --- |
| b. | Hyperthermia | d. | Long-bone fractures |

ANS:	C

Under normal conditions, proteins contribute only minor amounts to energy metabolism. Note that the percentage of protein used represents the normal turnover rate for replenishing structural and functional proteins in the body. Proteins might contribute significantly to energy expenditure, however, in cases of starvation.

PTS:	1	REF:	Page 276

103.	What percent change in metabolic rate would you expect if a patient has a fever of 103° F?

| a. | 10% increase | c. | 44% increase |
| --- | --- | --- | --- |
| b. | 20% increase | d. | 25% decrease |

ANS:	C

An increase in body temperature of 1° F will cause a 10% increase in metabolic rate. If the patient has a fever of 103° F, subtracting 98.6° F from the 103° F yields 4.4° F. A 1° F increase causes a 10% rise in metabolic rate, so 4.4  10% will equal a 44% increase.

PTS:	1	REF:	Page 276

104.	The expected respiratory quotient of a patient with severe sepsis is:

| a. | 0.7. | c. | 1.0. |
| --- | --- | --- | --- |
| b. | 0.8. | d. | >1.0. |

ANS:	A

Several studies have shown that patients with severe sepsis have respiratory quotients of approximately 0.7 because of their reliance on lipid metabolism for energy and an inability to use carbohydrates.

PTS:	1	REF:	Page 277

105.	In the below figure, A represents which of the following?

| a. | Vital capacity | c. | Total lung capacity |
| --- | --- | --- | --- |
| b. | Inspiratory capacity | d. | functional residual capacity |

ANS:	C

Total lung capacity is the sum of all the lung volumes.

PTS:	1	REF:	Page 255, Figure 8-19

106.	In the below figure, B represents which of the following?

| a. | Vital capacity | c. | Inspiratory reserve volume |
| --- | --- | --- | --- |
| b. | Inspiratory capacity | d. | Expiratory residual volume |

ANS:	A

Vital capacity is the total amount of air that can be exhaled after a maximum inspiration. It is the sum of the inspiratory reserve volume, the tidal volume, and the expiratory residual volume.

PTS:	1	REF:	Page 255, Figure 8-19

107.	In the below figure, C represents which of the following?

| a. | Vital capacity | c. | Functional residual capacity |
| --- | --- | --- | --- |
| b. | Inspiratory capacity | d. | Expiratory residual volume |

ANS:	B

Inspiratory capacity is the amount of air that can be inhaled after a normal exhalation.

PTS:	1	REF:	Page 255, Figure 8-19

108.	Which of the measurements shown in the figure is obtained by using body plethysmography?

| a. | B | c. | D |
| --- | --- | --- | --- |
| b. | C | d. | E |

ANS:	C

D represents the functional residual capacity and can be measured only by using either body plethysmography or inert gas techniques.

PTS:	1	REF:	Page 255, Figure 8-19

109.	In the below figure, A is pointing to which of the following components?

| a. | Dry seal | c. | CO2 absorber |
| --- | --- | --- | --- |
| b. | Water seal | d. | Linear potentiometer |

ANS:	C

A is pointing to the CO2 absorber.

PTS:	1	REF:	Page 239, Figure 8-2

110.	The device shown in the figure can measure which of the following?

| a. | FEV1 | c. | Forced vital capacity |
| --- | --- | --- | --- |
| b. | Minute volume | d. | Maximum expiratory flow |

ANS:	B

The figure shows a turbine flowmeter, handheld respirometer, which is used to measure tidal volume and minute volume.

PTS:	1	REF:	Page 242, Figure 8-6

111.	The figure represents which type of pneumotachograph?

| a. | Fleisch | c. | Vortex ultrasonic |
| --- | --- | --- | --- |
| b. | Variable-orifice | d. | Nonvortex ultrasonic |

ANS:	C

The figure represents a vortex ultrasonic pneumotachograph.

PTS:	1	REF:	Page 244, Figure 8-8

112.	In the figure, the amount of pressure required to overcome elastic forces is represented by which of the letters?

| a. | A | c. | C |
| --- | --- | --- | --- |
| b. | B | d. | D |

ANS:	C

Point C represents the Pplateau, which represents the total force that must be applied to overcome elastic forces only.

PTS:	1	REF:	Page 257, Figure 8-20

113.	The arrow points to which of the following lung volumes in the figure?

| a. | Tidal volume | c. | Inspiratory reserve volume |
| --- | --- | --- | --- |
| b. | Residual volume | d. | Expiratory residual volume |

ANS:	D

The arrow is pointing to expiratory residual volume, which is the total amount of gas that can be exhaled from the lung after a quiet exhalation.

PTS:	1	REF:	Page 255, Figure 8-19

114.	The capnogram seen in the figure shows which of the following?

| a. | Excessive phase 4 | c. | Indistinguishable phase 3 |
| --- | --- | --- | --- |
| b. | Deep “curare cleft” | d. | Slow cardiac oscillation |

ANS:	C

The figure shows an indistinguishable phase 3 caused by increases in physiologic dead space. This is a sign of chronic obstructive pulmonary disease.

PTS:	1	REF:	Page 269, Figure 8-31, A

115.	The capnogram seen in the figure shows which of the following?

| a. | Chronic obstructive pulmonary disease |
| --- | --- |
| b. | Cardiac oscillations |
| c. | Rebreathing exhaled gases |
| d. | Cheyne-Stokes breathing |

ANS:	A

Physiologic dead space increases with chronic obstructive pulmonary disease, which causes phase 3 to become indistinguishable, as shown in the figure.

PTS:	1	REF:	Page 269, Figure 8-31, E

116.	A decrease in perfusion of the alveoli will lead to which of the following?

| a. | Alveolar shunt | c. | Anatomic dead space |
| --- | --- | --- | --- |
| b. | Extra-alveolar shunt | d. | Physiologic dead space |

ANS:	D

A decrease of perfusion to the alveoli causes an increase in ventilation without perfusion, which is physiologic dead space. A shunt is perfusion without ventilation. Anatomic dead space is the normal amount of dead space in the body due to the conducting airways.

PTS:	1	REF:	Page 266, Figure 8-30

117.	The arterial-to-end-tidal CO2 and the arterial-to-maximum expiratory partial pressure of carbon dioxide (PCO2) gradients are both elevated and equal to each other. This finding is consistent with:

| a. | chronic obstructive pulmonary disease. |
| --- | --- |
| b. | left-sided heart failure. |
| c. | esophageal intubation. |
| d. | pulmonary embolism. |

ANS:	D

Normally, the difference between the PCO2 measured at the end of a maximum exhalation and the PaCO2 is minimal. Patients with chronic obstructive pulmonary disease or left-sided heart failure do not show an increase in this difference. However, patients with pulmonary embolism show an increased gradient.

PTS:	1	REF:	Page 270, Figure 8-32

118.	Phase 1 of a single-breath CO2 curve represents which of the following?

| a. | Anatomic dead space |
| --- | --- |
| b. | Physiologic dead space |
| c. | Transition between airway and alveolar exhaled gas |
| d. | Anatomic shunt |

ANS:	A

Anatomic dead space has no CO2; the mixed expired CO2 (FECO2) = 0 during that phase.

PTS:	1	REF:	Page 267

119.	In the figure, what does area X represent?

| a. | Amount of CO2 not eliminated because of alveolar dead space |
| --- | --- |
| b. | Amount of CO2 not eliminated due to anatomic dead space |
| c. | Amount of CO2 in a single breath |
| d. | Amount of CO2 rebreathed |

ANS:	C

Area X represents the actual amount of CO2 exhaled in the breath, assuming that no exhaled air is rebreathed. This is the volume of CO2 in a single breath.

PTS:	1	REF:	Page 270, Figure 8-34

120.	Which of the following statements is true if metabolism increases?

| a. | PaCO2 will rise even if ventilation increases. |
| --- | --- |
| b. | The single-breath CO2 curve (SBCO2) increases if ventilation does not. |
| c. | The volumetric CO2 tracing will not change even if ventilation increases. |
| d. | CO2 production decreases proportionally. |

ANS:	B

If metabolism increases and ventilation does not, PaCO2 rises and the amount of CO2 exhaled during the SBCO2 increases.

PTS:	1	REF:	Page 267