Chapter 12: Introduction to Ventilators

MULTIPLE CHOICE

1.	Pneumatic ventilators may incorporate which of the following components for operation?

1. Pistons

2. Air entrainers

3. Flexible diaphragms

4. Internal direct current batteries

| a. | 1 and 2 | c. | 1 and 4 |
| --- | --- | --- | --- |
| b. | 2 and 3 | d. | 3 and 4 |

ANS:	B

Pneumatic ventilators may incorporate components such as air entrainers, needle valves, flexible diaphragms, and spring-loaded valves to perform certain functions. Electrically powered ventilators most often use electrical power to run internal motors for operating pistons and may have internal direct current batteries, which can provide electrical power during patient transport or in the event of a power failure.

PTS:	1	REF:	Page 365

2.	An iron lung is which type of ventilator?

| a. | Fluidic | c. | Negative-pressure |
| --- | --- | --- | --- |
| b. | Positive-pressure | d. | Positive/negative pressure |

ANS:	C

An iron lung and a chest cuirass are both negative-pressure ventilators that are rarely used in the ventilation of patients these days.

PTS:	1	REF:	Page 366

3.	Which type of ventilator creates a sinusoidal pressure waveform pattern of positive and negative pressures?

| a. | Fluidic | c. | Positive-pressure |
| --- | --- | --- | --- |
| b. | Pneumatic | d. | Combined-pressure |

ANS:	D

A combined-pressure device, such as a high-frequency oscillator, produces oscillating gas pressure waveforms at the upper airway, producing a waveform that is sinusoidal, with positive-pressure and negative-pressure oscillations produced at the upper airway.

PTS:	1	REF:	Page 366

4.	An open-loop ventilator system can be described as which of the following?

| a. | Feedback | c. | Unintelligent |
| --- | --- | --- | --- |
| b. | Hierarchical | d. | Servo-controlled |

ANS:	C

Unintelligent ventilators are called open-loop systems because they do not respond to changes in the delivery of gas to the patient when there are any problems.

PTS:	1	REF:	Page 367

5.	A closed-loop ventilator system can be described as which of the following?

1. Feedback

2. Hierarchical

3. Unintelligent

4. Servo-controlled

| a. | 1 | c. | 1 and 4 |
| --- | --- | --- | --- |
| b. | 2 and 3 | d. | 2 and 4 |

ANS:	C

A closed-loop ventilator system is an intelligent system that can respond to certain inputs based on programming. This type of system is similar to cruise control on a car. When a parameter is set on the ventilator, the ventilator compares the input with the output and makes certain adjustments. These systems are also called feedback or servo-controlled systems.

PTS:	1	REF:	Page 367

6.	Which of the following is the flow waveform created by a linear-drive piston?

| a. | Sinusoidal | c. | Ascending ramp |
| --- | --- | --- | --- |
| b. | Rectangular | d. | Descending ramp |

ANS:	B

The rod of the piston moves the piston forward linearly inside a cylinder housing at a constant rate, producing a constant or rectangular waveform of gas flow to the patient.

PTS:	1	REF:	Page 371

7.	The function of the exhalation valve in the external circuit is to do which of the following?

| a. | Deflate during inspiration to allow the gas to go to the patient |
| --- | --- |
| b. | Inflate during inspiration to prevent leaks in the inspiratory circuit |
| c. | Inflate during inspiration to divert gas coming from the ventilator into the patient |
| d. | Inflate during expiration to allow the patient to exhale through the expiratory line |

ANS:	C

With an external circuit, gas flows both through the main inspiratory line to the patient and through the expiratory valve line during inspiration. This line inflates a balloon or puts pressure on a diaphragm, closing a hole. The hole in the exhalation valve is where the patient’s exhaled gases normally vent into the room. During inspiration, the hole is covered by the balloon or diaphragm. During exhalation, no gas goes through the main inspiratory line or the expiratory valve line. The balloon deflates, and the patient’s exhaled volume passes through the open hole.

PTS:	1	REF:	Page 369

8.	Which of the following flow-controlling valves operates by using the principles of electromagnetism?

| a. | Poppet | c. | Proportional solenoid |
| --- | --- | --- | --- |
| b. | Stepper motor | d. | Microprocessor-operated |

ANS:	C

Proportional solenoid valves operate on a basic principle of physics concerning electricity and magnetism.

PTS:	1	REF:	Page 373

9.	Which two of the following basic physical principles are used in the power transmission design of fluidic ventilators?

1. Venturi

2. Bernoulli

3. Beam deflection

4. Wall attachment

| a. | 1 and 2 | c. | 3 and 4 |
| --- | --- | --- | --- |
| b. | 2 and 3 | d. | 1 and 4 |

ANS:	C

Fluidic devices operate on two basic physical principles: wall attachment and beam deflection.

PTS:	1	REF:	Page 374

10.	A separation bubble in a fluidic device does which of the following?

| a. | Causes a flip-flop to occur |
| --- | --- |
| b. | Causes a high-pressure vortex |
| c. | Bends the gas jet stream toward the wall |
| d. | Pushes the gas jet stream away from the wall |

ANS:	C

A separation bubble is a low-pressure vortex that develops between a wall and a jet stream of gas. The bubble attracts or bends the jet stream toward the wall. The pocket of turbulence forms an air foil, similar to that seen with an airplane wing. When the gas entrained into the bubble from the jet stream equals the amount of air moving from the vortex flow of the bubble back to the jet stream, the attachment is stable.

PTS:	1	REF:	Page 375

11.	A patient receiving continuous positive airway pressure of 7.5 cm H2O via a mask and freestanding system appears to be in distress. The patient is using accessory muscles and is diaphoretic. The manometer is fluctuating between –5 cm H2O on inspiration and 7.5 cm H2O on expiration. The most apparent cause of the patient’s distress is which of the following?

| a. | Leak in the system | c. | Inadequate flow rate |
| --- | --- | --- | --- |
| b. | Improper mask fitting | d. | Obstruction of the threshold resistor |

ANS:	C

The patient is pulling in more than the flow of the continuous positive airway pressure system was set to deliver. This is apparent because of the wide fluctuation in the pressures. If there was a leak or improperly fitting mask, the pressure on expiration would not be at the set pressure. If there was an obstruction of the threshold resistor, it would prevent the patient’s exhaled gas from exiting and cause the pressure-release valve to pop off.

PTS:	1	REF:	Page 392

12.	The Downs continuous positive airway pressure system uses which of the following to establish positive airway pressure?

| a. | Weighted ball | c. | Spring-loaded valve |
| --- | --- | --- | --- |
| b. | Magnetic valve | d. | Opposing gas flow |

ANS:	C

The Downs continuous positive airway pressure generator uses a spring-loaded valve to adjust the amount of pressure.

PTS:	1	REF:	Page 377

13.	A patient is receiving continuous positive airway pressure from a freestanding, continuous-flow system. The respiratory therapist notices the pressure-release valve venting on each patient exhalation. What action should the respiratory therapist take?

| a. | Replace the continuous positive airway pressure device |
| --- | --- |
| b. | Decrease the flow rate |
| c. | Increase the pressure level |
| d. | Remove the pressure-release valve |

ANS:	A

When the pressure-release valve is activated, it indicates that the threshold resistor or spring-loaded valve is jammed, and the patient’s exhaled volume is not exiting through the continuous positive airway pressure valve. This is not proper function and indicates that the device needs to be replaced.

PTS:	1	REF:	Page 377

14.	A patient is receiving a continuous positive airway pressure of 10 cm H2O via mask from a Downs continuous positive airway pressure generator. During each inspiration, the pressure manometer decreases to 3 cm H2O. The patient is diaphoretic, tachypneic, and tachycardic. To correct this situation, the respiratory therapist should do which of the following?

| a. | Increase the gas flow | c. | Tighten the fit of the mask |
| --- | --- | --- | --- |
| b. | Increase the pressure | d. | Replace the spring-loaded valve |

ANS:	A

The patient is pulling in more than the flow of the continuous positive airway pressure system is able to provide. This is evident by the swing in pressures between 10 cm H2O and 3 cm H2O. Increasing the flow would decrease the patient’s work of breathing. Increasing pressure would make the patient work harder because of the inadequate flow rate. If there was a leak through an improperly fitting mask, very little pressure could be maintained. Replacement of the spring-loaded valve would not solve this problem.

PTS:	1	REF:	Page 377

15.	The set tidal volume is 500 mL. During the first breath, the exhaled volume measures 390 mL, with a peak pressure of 18 cm H2O. The second breath returns 460 mL, with 20 cm H2O. The third breath has a peak pressure of 22 cm H2O, with an exhaled volume of 500 mL. This can be described as what type of system?

| a. | Open-loop | c. | Single-circuit |
| --- | --- | --- | --- |
| b. | Closed-loop | d. | Double-circuit |

ANS:	B

This is a closed-loop system because the ventilator is making changes to the pressure based on the exhaled tidal volume. The pressure is being increased incrementally, which increases the exhaled volume to the set tidal volume.

PTS:	1	REF:	Page 367

16.	The gas that enters the ventilator from high-pressure sources goes directly to the patient. This is a description of a _____ system.

| a. | single-circuit | c. | pneumatic circuit |
| --- | --- | --- | --- |
| b. | double-circuit | d. | power transmission |

ANS:	A

In a single-circuit ventilator, the gas enters the ventilator and goes directly to the patient.

PTS:	1	REF:	Page 369

17.	Which of the following is a device that uses a flow resistor to create expiratory pressure?

| a. | Positive expiratory pressure mask |
| --- | --- |
| b. | Downs continuous positive airway pressure generator |
| c. | Modern ventilator |
| d. | Fluidic device |

ANS:	A

Flow resistors are used in positive expiratory pressure masks to combat atelectasis and aid in the removal of secretions. They are used less commonly in ventilator circuits. The Downs continuous positive airway pressure generator uses a threshold resistor to generate positive pressure. Modern ventilators use threshold resistors to generate expiratory pressure. Fluidic devices do not use flow resistors.

PTS:	1	REF:	Page 375

18.	The main parameter that the ventilator manipulates for any given breath is referred to as which of the following?

| a. | Control circuit | c. | Control variable |
| --- | --- | --- | --- |
| b. | Phase variable | d. | Cycle variable |

ANS:	C

The main parameter or variable controlled by the ventilator for any given breath is referred to as the control variable. A control circuit is part of the internal circuit of a mechanical ventilator and consists of mechanical devices, such as spring-loaded valves, pneumatic systems, or fluidic components. Phase variables are the variables that control the four phases of a breath. A cycle variable is the variable that is being used by the ventilator to end inspiration.

PTS:	1	REF:	Pages 380-381

19.	What variable begins the inspiratory phase of a ventilator?

| a. | Cycle | c. | Control |
| --- | --- | --- | --- |
| b. | Limit | d. | Trigger |

ANS:	D

The trigger variable begins inspiration. The cycle variable ends the inspiratory phase and begins exhalation. The limit is the maximum value for a variable that cannot be exceeded during inspiration. The control variable is one of the four elements of breath delivery that the ventilator is manipulating.

PTS:	1	REF:	Page 381

20.	Which of the following variables ends the inspiratory phase?

| a. | Cycle | c. | Trigger |
| --- | --- | --- | --- |
| b. | Limit | d. | Sensitivity |

ANS:	A

The cycle variable ends the inspiratory phase and begins exhalation.

PTS:	1	REF:	Page 381

21.	Which of the following variables begins a mandatory breath?

| a. | Flow | c. | Volume |
| --- | --- | --- | --- |
| b. | Time | d. | Pressure |

ANS:	B

Time-triggering the ventilator enables control of the beginning of inspiration based on the set mandatory rate. The breath is mandatory.

PTS:	1	REF:	Page 382

22.	The pressure maintained at the airway during exhalation is known as which of the following?

| a. | Baseline | c. | Sensitivity |
| --- | --- | --- | --- |
| b. | Proximal | d. | Mean airway |

ANS:	A

Baseline pressure is the pressure maintained at the airway during exhalation and the pressure from which inspiration begins. Proximal pressure is the pressure measured near the Y-connector of the ventilator circuit. Sensitivity is the control set by the operator that determines the amount of work the patient must do to trigger a breath. The mean airway pressure is the average pressure occurring in the airway during a complete respiratory cycle.

PTS:	1	REF:	Page 382

23.	Which two of the following patient triggers are most common?

1. Flow

2. Neural

3. Volume

4. Pressure

| a. | 1 and 2 | c. | 2 and 3 |
| --- | --- | --- | --- |
| b. | 1 and 4 | d. | 3 and 4 |

ANS:	B

Pressure and flow are the most common variables used for patient triggering. However, volume triggering and neural triggering also can be used.

PTS:	1	REF:	Page 381

24.	On a mechanical ventilator, pressure triggering is usually set within which of the following ranges?

| a. | –0.5 to –1.5 cm H2O | c. | 0.5 to 2.0 cm H2O |
| --- | --- | --- | --- |
| b. | –2.0 to 0.5 cm H2O | d. | 1.0 to 3.0 cm H2O |

ANS:	A

Pressure triggering usually is set at –0.5 to –1.5 cm H2O. This means that the trigger will occur at 0.5 to 1.5 cm H2O below the baseline pressure.

PTS:	1	REF:	Page 382

25.	If the base flow is 8 L/min and the flow-trigger setting is 3 L/min, at what flow rate would the ventilator be triggered?

| a. | 3 L/min | c. | 8 L/min |
| --- | --- | --- | --- |
| b. | 5 L/min | d. | 11 L/min |

ANS:	B

The ventilator will begin inspiration when the expiratory flow is measured at 5 L/min (8 L/min – 3 L/min = 5 L/min).

PTS:	1	REF:	Page 382

26.	If the base flow is 10 L/min and the flow-trigger setting is 4 L/min, the flow during the expiratory phase is _____ L/min.

| a. | 4 | c. | 10 |
| --- | --- | --- | --- |
| b. | 6 | d. | 14 |

ANS:	C

The base flow is the bias or background flow that is present during the expiratory phase. That flow is the set base flow, or 10 L/min in this case.

PTS:	1	REF:	Page 382

27.	With a baseline pressure setting of 6 cm H2O and a sensitivity setting of –1.5 cm H2O, inspiration will begin at _____ cm H2O.

| a. | –1.5 | c. | +1.5 |
| --- | --- | --- | --- |
| b. | –4.5 | d. | +4.5 |

ANS:	D

The actual pressure at which the breath will be triggered is +4.5 cm H2O (6 cm H2O – 1.5 cm H2O = 4.5 cm H2O).

PTS:	1	REF:	Page 382

28.	What is the flow rate if the set volume is 550 mL and the inspiratory time is 1.2 second?

| a. | 27.5 L/min | c. | 55 L/min |
| --- | --- | --- | --- |
| b. | 45.8 L/min | d. | 66 L/min |

ANS:	A

Volume = Flow  Time, so Flow = Volume ÷ Time. 550 mL = 0.55 L. Flow = 0.55 L ÷ 1.2 s = 0.46 L/s  60 = 27.5 L/min.

PTS:	1	REF:	Page 380

29.	Decreased lung compliance during pressure-limited ventilation would cause which of the following?

| a. | Increased flow rates | c. | Decreased volume delivery |
| --- | --- | --- | --- |
| b. | Increased pressure delivery | d. | Decreased inspiratory pressures |

ANS:	C

During pressure-targeted ventilation, volume delivery varies with the set inspiratory time and the patient’s lung characteristics. When the patient’s lung compliance is decreased, the lungs are stiff, and less volume is delivered for the same amount of pressure.

PTS:	1	REF:	Page 380

30.	Which of the following variables cannot be exceeded during inspiration?

| a. | Limit | c. | Safety |
| --- | --- | --- | --- |
| b. | Cycle | d. | Trigger |

ANS:	A

A limiting variable has a maximum value that cannot be exceeded during inspiration because the ventilator does not allow it; however, reaching the limit does not end inspiratory flow to the patient.

PTS:	1	REF:	Page 386

31.	The maximum safety pressure should be set at _____ cm H2O when the peak pressure reached during volume-controlled ventilation is 25 cm H2O.

| a. | 30 | c. | 40 |
| --- | --- | --- | --- |
| b. | 35 | d. | 45 |

ANS:	B

The maximum safety pressure is usually set by the operator at 10 cm H2O above the peak pressure reached during inspiration.

PTS:	1	REF:	Page 387

32.	When the maximum safety pressure setting on an adult ventilator is reached, which of the following occurs?

| a. | Inspiration is triggered. |
| --- | --- |
| b. | Flow begins to decrease. |
| c. | Inspiration ends prematurely. |
| d. | Pressure is held until the target volume is delivered. |

ANS:	C

When the maximum safety pressure is reached, the ventilator will sound an alarm, limit the pressure, and cycle out of inspiration, ending inspiration early.

PTS:	1	REF:	Page 388

33.	When the maximum safety pressure on a volume ventilator is reached, which of the following occurs?

| a. | The flow rate begins to taper. |
| --- | --- |
| b. | Pressure is held until the time cycle is reached. |
| c. | Pressure is held until the ventilator volume cycles. |
| d. | Inspiration ends before the entire volume is delivered. |

ANS:	D

When the maximum safety pressure is reached, the ventilator sounds an alarm, limits the pressure, and cycles out of inspiration, ending inspiration early.

PTS:	1	REF:	Page 388

34.	Which of the following represents the amount of time for a pressure-targeted ventilator breath to reach the set pressure?

| a. | Rise time | c. | Total cycle time |
| --- | --- | --- | --- |
| b. | Inspiratory time | d. | Inspiratory hold |

ANS:	B

Inspiratory time dictates the amount of time that a ventilator spends in the inspiratory phase.

PTS:	1	REF:	Page 366, Figure 12-2

35.	The flow, volume, and pressure waveforms characteristic of a pressure-limited, time-cycled breath are represented in which column of the figure?

| a. | A | c. | C |
| --- | --- | --- | --- |
| b. | B | d. | D |

ANS:	C

The pressure waveform shown in column C has the characteristic flat top of a pressure limit. Because the breath does not end when the pressure limit is met, the breath is time cycled.

PTS:	1	REF:	Page 391, Figure 12-38

36.	The flow, volume, and pressure waveforms characteristic of volume breaths with a constant flow are shown in which column of the figure?

| a. | A | c. | C |
| --- | --- | --- | --- |
| b. | B | d. | D |

ANS:	D

A constant-flow breath creates a rectangular or square flow waveform, as seen in column D.

PTS:	1	REF:	Page 391, Figure 12-38

37.	The flow, volume, and pressure waveforms characteristic of volume breaths with a descending ramp flow pattern are shown in which column of the figure?

| a. | A | c. | C |
| --- | --- | --- | --- |
| b. | B | d. | D |

ANS:	A

The volume breath shown in column A has a descending ramp flow waveform.

PTS:	1	REF:	Page 391, Figure 12-38

38.	The flow, volume, and pressure waveforms characteristic of volume breaths with a sine-like flow pattern are shown in which column of the figure?

| a. | A | c. | C |
| --- | --- | --- | --- |
| b. | B | d. | D |

ANS:	B

The sine-like flow waveform begins at zero and reaches the maximum inspiratory flow in the middle of inspiration. It then tapers down to zero.

PTS:	1	REF:	Page 391, Figure 12-48

39.	The figure shows the pressure-time scalar for a pressure-support breath. To alleviate the problem depicted in this figure, the respiratory therapist should do which of the following?

| a. | Decrease the pressure setting | c. | Add inspiratory sloping |
| --- | --- | --- | --- |
| b. | Increase the inspiratory time | d. | Decrease the set volume |

ANS:	C

The scalar shows a pressure overshoot owing to the rapid flow rate at the very beginning of inspiration. Adjusting the slope of the pressure and flow curves can change the rate of gas delivery at the very beginning of inspiration. This tapers the curves slightly until they reach their maximum value. Decreasing the pressure setting may alleviate the overshooting problem, but it does not provide the same level of assistance to the patient. Inspiratory time and volume are controlled by the patient in pressure-support ventilation and cannot be adjusted.

PTS:	1	REF:	Page 400

40.	Improving lung compliance during volume-controlled ventilation would cause a(n) _____ in peak _____.

| a. | decease; pressure | c. | decrease; flow rate |
| --- | --- | --- | --- |
| b. | increase; pressure | d. | increase; flow rate |

ANS:	A

During volume-controlled ventilation, the peak inspiratory pressure generated by the delivery of the volume varies according to the characteristics of the lungs. Improving compliance makes it easier to deliver a volume, decreasing the pressure generated in the lungs.

PTS:	1	REF:	Page 398

41.	Worsening lung compliance during volume-controlled ventilation would cause which of the following?

| a. | A decrease in peak pressure | c. | A decrease in volume delivery |
| --- | --- | --- | --- |
| b. | An increase in peak pressure | d. | An increase in peak flow rate |

ANS:	B

During volume-controlled ventilation, the peak inspiratory pressure generated by the delivery of the volume varies according to the characteristics of the lungs. Worsening lung compliance is indicative of stiffening lungs. This would cause more pressure to be generated during the delivery of a set volume.

PTS:	1	REF:	Page 398

42.	Improving lung compliance during pressure-controlled ventilation would cause which of the following?

| a. | A decrease in peak pressure | c. | An increase in volume delivery |
| --- | --- | --- | --- |
| b. | An increase in peak pressure | d. | A decrease in volume delivery |

ANS:	C

During pressure-controlled ventilation, the delivery of the volume varies according to the characteristics of the lungs. Improving compliance makes it easier to deliver a volume with the same set pressure.

PTS:	1	REF:	Page 398

43.	Worsening lung compliance during pressure-controlled ventilation would cause which of the following?

| a. | A decrease in peak pressure | c. | A decrease in volume delivery |
| --- | --- | --- | --- |
| b. | An increase in peak pressure | d. | An increase in peak flow rate |

ANS:	C

During pressure-controlled ventilation, the delivery of the volume varies according to the characteristics of the lungs. Worsening lung compliance is indicative of stiffening lungs. This would cause less volume to be delivered with the same set pressure.

PTS:	1	REF:	Page 398

44.	During pressure-targeted ventilation with a set pressure of 24 cm H2O, the patient coughs. The ventilator responds by in which of the following ways?

| a. | Inspiration ends immediately. |
| --- | --- |
| b. | Inspiratory time increases. |
| c. | Inspiration continues with the set pressure. |
| d. | Inspiratory pressure ends at the maximum safety pressure. |

ANS:	D

It is a common misconception that the pressure-target setting cannot be exceeded during pressure-targeted ventilation. A cough by the patient can cause the pressure to rise above the set value. This is why it is vital that a safe upper pressure limit, or maximum safety pressure, be set at approximately 10 cm H2O above the pressure-target setting during pressure ventilation.

PTS:	1	REF:	Page 388

45.	What is the compressibility factor for the patient circuit when the ventilator volume is set at 200 mL, the recorded peak inspiratory pressure is 72 cm H2O, and the measured volume is 148 mL?

| a. | 0.36 cm H2O/mL | c. | 2.06 mL/cm H2O |
| --- | --- | --- | --- |
| b. | 0.49 cm H2O/mL | d. | 2.78 mL/cm H2O |

ANS:	C

The compressibility factor of a patient circuit is the circuit’s tubing compliance. Tubing compliance equals the measured volume divided by the measured peak inspiratory pressure (Tubing compliance = Volume/Peak inspiratory pressure). Tubing compliance = 148 mL ÷ 72 cm H2O = 2.06 mL/cm H2O.

PTS:	1	REF:	Page 392, Box 12-12

46.	What is the compressibility factor for the patient circuit when the ventilator volume is set at 100 mL, the recorded peak inspiratory pressure is 41 cm H2O, and the measured volume is 68 mL?

| a. | 0.41 cm H2O/mL | c. | 1.66 mL/cm H2O |
| --- | --- | --- | --- |
| b. | 0.60 cm H2O/mL | d. | 2.44 mL/cm H2O |

ANS:	C

The compressibility factor of a patient circuit is the circuit’s tubing compliance. Tubing compliance equals the measured volume divided by the measured peak inspiratory pressure (Tubing compliance = Volume/Peak inspiratory pressure). Tubing compliance = 68 mL ÷ 41 cm H2O = 1.66 mL/cm H2O.

PTS:	1	REF:	Page 392, Box 12-12

47.	What is the volume that will reach the patient when the volume setting is 700 mL, the tubing compliance is 2 mL/cm H2O, and the measured peak inspiratory pressure is 35 cm H2O?

| a. | 770 mL | c. | 665 mL |
| --- | --- | --- | --- |
| b. | 698 mL | d. | 630 mL |

ANS:	D

Volume lost because of tubing compliance is equal to the set volume minus the tubing compliance multiplied by the measured peak inspiratory pressure. Volume lost = 700 mL ?2- (2 mL/cm H2O  35 cm H2O) = 700 mL ?2- 70 mL = 630 mL.

PTS:	1	REF:	Page 390

48.	How much volume is lost in the ventilator circuit when the volume is set at 500 mL, the peak inspiratory pressure is 40 cm H2O, and the tubing compliance is 3 mL/cm H2O?

| a. | 380 mL | c. | 150 mL |
| --- | --- | --- | --- |
| b. | 280 mL | d. | 120 mL |

ANS:	D

Volume lost in the ventilator circuit equals tubing compliance multiplied by the measured peak inspiratory pressure. Volume lost = 3 mL/cm H2O  40 cm H2O = 120 mL.

PTS:	1	REF:	Page 390

49.	Pressure-support ventilation is classified as:

1. Flow cycled

2. Time triggered

3. Volume targeted

4. Pressure targeted

5. Patient triggered

| a. | 2 and 3 | c. | 1, 4, and 5 |
| --- | --- | --- | --- |
| b. | 4 and 5 | d. | 1, 2, and 3 |

ANS:	C

In pressure-support ventilation, when the patient begins inspiration, inspiration is limited by pressure and ends when the patient’s peak inspiratory flow decreases by a certain percentage (flow cycled).

PTS:	1	REF:	Page 387

50.	Which of the following flow-termination settings for a pressure-supported breath that has a peak inspiratory flow of 30 L/min would allow enough time for a visible plateau on the pressure-time curve?

| a. | 5 L/min | c. | 20% of peak flow |
| --- | --- | --- | --- |
| b. | 20 L/min | d. | 30% of peak flow |

ANS:	A

The longest inspiratory time provides a visible plateau on the pressure-time scalar. This is caused by a flow termination that takes the longest time to reach. The flow termination that is lowest provides the most visible plateau on a pressure-time scalar. Because 20% of the peak flow of 30 L/min is 6 L/min and 30% of 30 L/min is 9 L/min, the lowest flow in this case is 5 L/min.

PTS:	1	REF:	Page 390

51.	Plateau pressure may be obtained during volume ventilation by using which of the following?

| a. | Expiratory pause | c. | End-expiratory pressure |
| --- | --- | --- | --- |
| b. | Inspiratory pause | d. | Increased inspiratory time |

ANS:	B

During volume ventilation, plateau pressure is obtained by using the maneuver that extends or delays exhalation by keeping the expiratory valve momentarily closed. This prevents the gas from leaving the circuit. The maneuver is known as an inspiratory pause, inspiratory hold, or inspiratory plateau.

PTS:	1	REF:	Page 390

52.	During volume ventilation, peak inspiratory pressure is the pressure needed to overcome which two of the following factors?

1. Muscle pressure

2. Airflow resistance

3. Elastic recoil pressure

4. Transthoracic pressure

| a. | 1 and 2 | c. | 3 and 4 |
| --- | --- | --- | --- |
| b. | 1 and 3 | d. | 2 and 3 |

ANS:	A

During volume ventilation, pressure increases as the volume is being delivered. The pressure at the end of inspiration, or peak inspiratory pressure, represents the pressure needed to overcome both airflow resistance and the elastic recoil pressure of the lungs and thorax. This concept is demonstrated in the equation of motion: Muscle pressure + Ventilator pressure = Elastic recoil pressure + Flow resistance pressure. Muscle pressure is the pressure generated by a person’s muscles of ventilation to overcome elastic recoil pressure and airflow resistance. Transthoracic pressure is a pressure gradient between the pleural space and the body surface. During spontaneous breathing, it is the total pressure needed to expand or contract the lungs and chest wall together.

PTS:	1	REF:	Page 391, Box 12-11

53.	The transairway pressure represents the pressure associated with which of the following?

| a. | Elastic recoil | c. | Tubing compressibility |
| --- | --- | --- | --- |
| b. | Airflow resistance | d. | End-expiratory pressure |

ANS:	B

The transairway pressure is the pressure required to overcome airflow resistance and is calculated by subtracting the plateau pressure from the peak inspiratory pressure.

PTS:	1	REF:	Page 391, Box 12-11

54.	Increased transairway pressure is associated with which two of the following changes?

1. Increased airway resistance

2. Decreased airway resistance

3. Increased static compliance

4. Decreased static compliance

| a. | 3 and 4 | c. | 2 and 4 |
| --- | --- | --- | --- |
| b. | 1 and 3 | d. | 1 and 2 |

ANS:	B

Transairway pressure is the pressure required to overcome airflow resistance and is calculated by subtracting the plateau pressure from the peak inspiratory pressure. An increase in the peak inspiratory pressure with a constant plateau pressure increases the transairway pressure. Another way to increase transairway pressure would be to decrease the plateau pressure (i.e., improved static compliance) and keep a constant peak inspiratory pressure.

PTS:	1	REF:	Page 391, Box 12-11

55.	Which of the following is the formula for transairway pressure?

| a. | Peak inspiratory pressure – Plateau pressure |
| --- | --- |
| b. | Plateau pressure – Peak inspiratory pressure |
| c. | Positive end-expiratory pressure – Plateau pressure |
| d. | Peak inspiratory pressure – Positive end-expiratory pressure |

ANS:	A

The transairway pressure is the pressure required to overcome airflow resistance and is calculated by subtracting the plateau pressure from the peak inspiratory pressure.

PTS:	1	REF:	Page 392, Box 12-11

56.	What is the transairway pressure from the information found in the figure?

| a. | 10 cm H2O | c. | 25 cm H2O |
| --- | --- | --- | --- |
| b. | 15 cm H2O | d. | 35 cm H2O |

ANS:	B

Transairway pressure = Peak inspiratory pressure – Plateau pressure = 35 cm H2O – 20 cm H2O = 15 cm H2O.

PTS:	1	REF:	Page 391, Box 12-11

57.	What is the static lung compliance when the returned tidal volume is 500 mL, the baseline pressure is zero, and the plateau pressure reading is 25 cm H2O?

| a. | 0.02 L/cm H2O | c. | 0.2 L/cm H2O |
| --- | --- | --- | --- |
| b. | 2.0 L/cm H2O | d. | 2.0 L/cm H2O |

ANS:	A

Static lung compliance = Volume ÷ (Plateau pressure – Positive end-expiratory pressure) = 0.5 L ÷ (25 cm H2O – 0 cm H2O) = 0.02 L/cm H2O.

PTS:	1	REF:	Page 390

58.	What is the static lung compliance when the returned tidal volume is 700 mL, the positive end-expiratory pressure is 10 cm H2O, the peak inspiratory pressure is 45 cm H2O, and the plateau pressure is 35 cm H2O?

| a. | 0.160 L/cm H2O | c. | 0.028 L/cm H2O |
| --- | --- | --- | --- |
| b. | 0.020 L/cm H2O | d. | 0.735 L/cm H2O |

ANS:	B

Static lung compliance = Volume ÷ (Plateau pressure – Positive end-expiratory pressure) = 0.7 L ÷ (45 cm H2O – 10 cm H2O) = 0.7 L ÷ 35 cm H2O = 0.02 L/cm H2O.

PTS:	1	REF:	Page 390

59.	What is the static lung compliance when the returned tidal volume is 460 mL, the positive end-expiratory pressure is 8 cm H2O, the peak inspiratory pressure is 27 cm H2O, and the plateau pressure is 18 cm H2O?

| a. | 20 mL/cm H2O | c. | 46 mL/cm H2O |
| --- | --- | --- | --- |
| b. | 26 mL/cm H2O | d. | 51 mL/cm H2O |

ANS:	B

Static lung compliance = Volume ÷ (Plateau pressure – Positive end-expiratory pressure) = 460 mL ÷ (18 cm H2O – 8 cm H2O) = 460 mL ÷ 10 cm H2O = 46 mL/cm H2O.

PTS:	1	REF:	Page 390

60.	Increasing the functional residual capacity of a spontaneously breathing patient to improve oxygenation can be accomplished by which of the following?

| a. | Inspiratory plateau | c. | Negative end-expiratory pressure |
| --- | --- | --- | --- |
| b. | Positive end-expiratory pressure | d. | Continuous positive airway pressure |

ANS:	D

Continuous positive airway pressure is a technique in which a patient breathes spontaneously at an elevated baseline pressure. This increases the functional residual capacity and mean airway pressure and improves lung recruitment and oxygenation.

PTS:	1	REF:	Page 392

61.	What procedure is used to estimate the amount of air trapped in the patient’s lungs and in the patient circuit?

| a. | Inspiratory pause | c. | Pressure-control ventilation |
| --- | --- | --- | --- |
| b. | End-expiratory pause | d. | Pressure-release ventilation |

ANS:	B

End-expiratory pause, or expiratory hold, is a procedure performed to estimate the pressure in the patient’s lung and ventilator circuit caused by trapped air.

PTS:	1	REF:	Page 392

62.	Auto-PEEP can be detected without the use of an expiratory pause on which of the following curves?

| a. | Pressure-time | c. | Pressure-volume |
| --- | --- | --- | --- |
| b. | Flow-time | d. | Volume-time |

ANS:	B

When a patient has auto-PEEP, the expiratory flow does not return to zero. A new inspiration begins before the patient has had time to exhale completely. As a result, air remains trapped in the lungs. A pressure-time curve shows auto-PEEP only when an expiratory pause is used.

PTS:	1	REF:	Page 393

63.	What does point A represent on the pressure-time scalar shown in the figure?

| a. | Plateau pressure | c. | Positive end-expiratory pressure |
| --- | --- | --- | --- |
| b. | Maximum safety pressure | d. | Auto-PEEP |

ANS:	D

Auto-PEEP is made visible on a pressure-time scalar with the use of an end-expiratory pause. Point A is not positive end-expiratory pressure because there are areas in the waveform that go down to zero.

PTS:	1	REF:	Page 401, Figure 12-48

64.	Which of the following is the technique that can actively remove gas from a ventilator circuit at the beginning of expiration?

| a. | Expiratory pause | c. | Positive end-expiratory pressure |
| --- | --- | --- | --- |
| b. | Inspiratory pause | d. | Negative end-expiratory pressure |

ANS:	D

If negative pressure is applied on exhalation, the decrease in pressure draws air out of the ventilator circuit. An expiratory pause facilitates the measurement of auto-PEEP. An inspiratory pause facilitates the calculation of static pressure and transairway pressure. Positive end-expiratory pressure is positive pressure applied on exhalation and facilitates the improvement of oxygenation.

PTS:	1	REF:	Page 392

65.	Which of the following statements is true concerning mechanical ventilation classification?

| a. | The names for modes of ventilation are current and accurate. |
| --- | --- |
| b. | Mode classification should be individualized by manufacturers. |
| c. | Mode classification terminology needs to be clearly defined. |
| d. | All manufacturer mode terminology adheres to Chatburn’s system. |

ANS:	C

After years of studying the engineering, clinical aspects, and resources needed to train clinicians about ventilation modes, Chatburn reached four main conclusions about ventilator modes: (1) the current names for modes of ventilation are outdated and confusing, (2) confusion about names can lead to confusion about clinical application, (3) a ventilator mode classification system should be applicable to the modes available on any ventilator, and (4) all terminology proposed for ventilator classification needs to be clearly defined.

PTS:	1	REF:	Page 396

66.	Which of the following is a ventilator mode that allows the ventilator to switch from pressure control to volume control or vice versa?

| a. | Dual-control mode | c. | Volume-control mode |
| --- | --- | --- | --- |
| b. | Flow-control mode | d. | Pressure-control mode |

ANS:	A

A dual-control, breath-variable mode allows the ventilator to switch from pressure control to volume control or vice versa within a breath. In this type of mode, the primary breath control is both volume and pressure.

PTS:	1	REF:	Page 396

67.	A dual mode that features a pressure-controlled breath pattern with volume as a conditional variable is which of the following?

| a. | Tactical control ventilation | c. | Pressure-regulated volume control |
| --- | --- | --- | --- |
| b. | Pressure-targeted ventilation | d. | Volume-regulated pressure control |

ANS:	C

Pressure-regulated volume control is a breath pattern for which the pressure-limit value is set by the operator, along with the minimum volume that needs to be delivered. The conditional variable, volume, is measured by the ventilator, and the pressure setting may be increased or decreased depending on that measurement.

PTS:	1	REF:	Page 396

68.	Which of the following ventilator control types consists of the output of the ventilator automatically matching a constant operator-preset input value?

| a. | Optimal | c. | Auto-set-point |
| --- | --- | --- | --- |
| b. | Set-point | d. | Knowledge-based |

ANS:	B

A set-point control-type example mode is available in volume-controlled intermittent volume ventilation; in this mode, mandatory breaths are volume limited and volume cycled for the tidal volume setting. Optimal control has one ventilator set-point automatically adjusted to optimize another set-point according to some model of system behavior, and the output can be maximized or minimized dynamically. The auto-set-point control type enables the ventilator to select which operator-adjusted set-points are enforced at the moment. The knowledge-based control type has set-points that are automatically adjusted according to a rule-based expert system.

PTS:	1	REF:	Page 396, Table 12-1

69.	Which of the following control modes of ventilation consists of pressure-limited mandatory breaths that can be automatically adjusted between breaths to achieve a preset tidal volume?

| a. | Servo | c. | Adaptive |
| --- | --- | --- | --- |
| b. | Optimal | d. | Auto-set-point |

ANS:	C

An example of an adaptive control type is pressure-regulated volume control, in which the mandatory breaths are pressure limited and can be adjusted between breaths to achieve a preset tidal volume chosen by the operator. In the servo-controlled mode, the ventilator output automatically follows a varying input by the patient. In the optimal control type, one ventilator set-point is automatically adjusted to optimize another set-point according to some model of system behavior, and the output can be maximized or minimized dynamically. The auto-set-point control type allows the ventilator to select which operator-adjusted set-point to enforce at the moment.

PTS:	1	REF:	Page 396, Table 12-1

70.	Which of the following control modes of ventilation consists of ventilator output automatically following a varying input (e.g., volume or flow generated by the patient)?

| a. | Servo | c. | Neural network |
| --- | --- | --- | --- |
| b. | Optimal | d. | Knowledge-based |

ANS:	A

An example of a servo-controlled mode is the proportional assist mode, in which the instantaneous value of pressure is proportional to the instantaneous volume or flow generated by the patient. The optimal control type has one ventilator set-point automatically adjusted to optimize another set-point according to some model of system behavior, and the output can be maximized or minimized dynamically. The artificial neural network is based on acceptable arterial blood gases to maintain patient ventilation (experimental). The knowledge-based control type has set-points that are automatically adjusted according to a rule-based expert system such as end-tidal carbon dioxide.

PTS:	1	REF:	Page 396, Table 12-1

71.	The mandatory breath begins as a pressure-controlled breath, but it switches to volume control during inspiration because a preset volume is not achieved. This describes which of the following control types?

| a. | Set-point | c. | Auto-set-point |
| --- | --- | --- | --- |
| b. | Adaptive | d. | Knowledge-based |

ANS:	C

In this example, inspiration is starting out in pressure control and then switches to volume control by using operator-adjusted set-points. One type of mode that fits this category of control types is pressure augmentation (BEAR 1000). In the set-point control type, the output of the ventilator automatically matches a constant operator-preset input value. The adaptive control type changes set-point types based on the patient’s changing condition. The knowledge-based control type has the set-points automatically adjusted according to a rule-based expert system, such as breathing frequency, tidal volume, and end-tidal carbon dioxide levels.

PTS:	1	REF:	Page 396, Table 12-1

72.	The pressure support level is automatically adjusted by the ventilator to maintain appropriate breathing frequency, tidal volume, and end-tidal carbon dioxide, depending on the patient’s pathologic condition. This mode description pertains to which of the following ventilator control types?

| a. | Adaptive | c. | Knowledge-based |
| --- | --- | --- | --- |
| b. | Set-point | d. | Neurally adjusted ventilatory assist |

ANS:	C

The knowledge-based control type has the set-points automatically adjusted according to a rule-based expert system, such as breathing frequency, tidal volume, and end-tidal carbon dioxide levels. An example of this kind of control type is SmartCare on the Dräger EvitaXL. The adaptive control type changes set-point types based on the patient’s changing condition. For example, mandatory breaths are pressure limited, and this pressure-limit setting is automatically adjusted between breaths to achieve the preset tidal volume. In the set-point control type, the output of the ventilator automatically matches a constant operator-preset input value. Neurally adjusted ventilatory assist is not a control type. It is a mode that adjusts set-points automatically, based on the electrical activity of the diaphragm and the level of response to the electrical activity as set by the operator.

PTS:	1	REF:	Page 396, Table 12-1

73.	The description “a mode where each breath is pressure limited and the pressure-limit level is adjusted automatically between breaths by using ventilatory mechanics to minimize the patient’s work of breathing” pertains to which ventilator control type?

| a. | Servo | c. | Adaptive |
| --- | --- | --- | --- |
| b. | Optimal | d. | Auto-set-point |

ANS:	B

The optimal control type has one ventilator set-point—in this case, pressure limit—that automatically adjusts to optimize another set-point, which in this case is the patient’s ventilatory mechanics, according to some model of system behavior, and the output can be maximized or minimized dynamically. With the servo-controlled type, the ventilator output automatically follows a varying input from the patient. The adaptive control type automatically adjusts one set-point to achieve another set-point as the patient’s condition changes, such as pressure-limit changing to achieve a preset tidal volume. The auto-set-point control type enables the selection of which operator-adjusted set-point to enforce at the moment.

PTS:	1	REF:	Page 396, Table 12-1

74.	The ventilator mode that is either patient triggered or time triggered and allows the patient to breathe spontaneously between mandatory ventilator breaths is known as:

| a. | Pressure-support ventilation | c. | Intermittent mandatory ventilation |
| --- | --- | --- | --- |
| b. | Pressure-controlled ventilation | d. | Continuous mandatory ventilation |

ANS:	C

Intermittent mandatory ventilation is designed to deliver volume-targeted or pressure-targeted breaths at a set minimum frequency (i.e., it is time triggered). Between mandatory breaths, the patient can breathe spontaneously from the ventilator circuit without getting the set volume or pressure from the ventilator itself.

PTS:	1	REF:	Page 395

75.	Mandatory breaths during pressure-controlled mechanical ventilation (pressure-controlled continuous mandatory ventilation) are cycled into expiration by which of the following?

| a. | Time | c. | Volume |
| --- | --- | --- | --- |
| b. | Flow | d. | Pressure |

ANS:	A

All pressure-controlled continuous mandatory ventilation breaths are time cycled.

PTS:	1	REF:	Page 385

76.	Which of the following ventilator modes decreases the effect of airway resistance on the patient’s spontaneous breathing?

| a. | Pressure-support ventilation |
| --- | --- |
| b. | Pressure-controlled continuous mandatory ventilation |
| c. | Volume-controlled intermittent mandatory ventilation |
| d. | Volume-controlled continuous mandatory ventilation |

ANS:	A

Pressure-support ventilation is a spontaneous breath type of ventilation that allows the operator to select a pressure to support the patient’s work of breathing and counteract airway resistance.

PTS:	1	REF:	Page 407

77.	Which of the following represents the safe starting point for pressure support of a patient just placed on volume-controlled intermittent mandatory ventilation with the following set and measured parameters: tidal volume, 475 mL; peak inspiratory pressure, 30 cm H2O; plateau pressure, 10 cm H2O; and flow rate, 45 L/min?

| a. | 4 cm H2O | c. | 20 cm H2O |
| --- | --- | --- | --- |
| b. | 7 cm H2O | d. | 27 cm H2O |

ANS:	C

The safe starting point for pressure support is the transairway pressure value (Peak inspiratory pressure – Plateau pressure). In this case, transairway pressure = 30 cm H2O – 10 cm H2O = 20 cm H2O.

PTS:	1	REF:	Page 408

78.	Adult candidates for pressure-support ventilation include patients with which of the following?

1. Nasotracheal tube

2. Minute ventilation <10 L

3. Spontaneous tidal volume <100 mL

4. Spontaneous respiratory rate >20 breaths/min

| a. | 1 and 2 | c. | 2, 3, and 4 |
| --- | --- | --- | --- |
| b. | 1 and 4 | d. | 1, 3, and 4 |

ANS:	D

Patients who have nasotracheal tubes usually have tubes that are smaller than they would be if they were intubated orally. This increases the resistance through the artificial airway, making pressure support necessary when the patient is breathing spontaneously. Patients with small spontaneous tidal volumes can be assisted with pressure support to increase their tidal volumes. The pressure-support ventilation level can be adjusted to attain the desired tidal volume based on the patient’s ideal body weight. Patients with high respiratory rates can also benefit from pressure-support ventilation. Titrating the pressure support level to the patient’s respiratory rate and tidal volume provides the appropriate assistance to overcome the airway and artificial tubing resistance and improve the patient’s work of breathing.

PTS:	1	REF:	Page 409, Box 12-17

79.	Which of the following ventilator modes is recommended for patients who have intact respiratory centers and who are breathing spontaneously?

| a. | Intermittent mandatory ventilation | c. | Pressure-controlled ventilation |
| --- | --- | --- | --- |
| b. | Pressure-support ventilation | d. | Continuous mandatory ventilation |

ANS:	B

Pressure-support ventilation is a spontaneous mode of ventilation, and it does not provide ventilator breaths to the patient. The patient must have an intact respiratory center and must be able to breathe spontaneously.

PTS:	1	REF:	Page 408

80.	Which of the following is a mode of ventilation that is time triggered or patient triggered as well as pressure limited and time cycled?

| a. | Pressure-support ventilation |
| --- | --- |
| b. | Pressure-controlled intermittent mandatory ventilation |
| c. | Pressure-controlled continuous mandatory ventilation |
| d. | Dual control continuous mandatory ventilation |

ANS:	C

Time triggered means that there are ventilator-initiated breaths. Patient triggered means that the patient is allowed to do the work to trigger a ventilator breath. Each breath is ended by a set inspiratory time. This is the description of the pressure-controlled continuous mandatory ventilation mode. All pressure-support ventilation breaths are patient triggered and flow cycled. Pressure-controlled intermittent mandatory ventilation breaths allow spontaneous breaths between ventilator breaths, and those breaths are flow cycled. Dual control continuous mandatory ventilation is a dual control mode of ventilation in which the breath may start out as pressure-limited but switch to volume-limited if the volume setting is not reached.

PTS:	1	REF:	Page 396

81.	What is the mode depicted in the pressure-versus-time waveform in the figure?

| a. | Airway pressure-release ventilation |
| --- | --- |
| b. | Bilevel positive airway pressure |
| c. | Pressure-controlled intermittent mandatory ventilation plus pressure support |
| d. | Volume-controlled intermittent mandatory ventilation plus pressure support |

ANS:	C

The larger “flat-top” pressure curves represent patient-triggered, pressure-limited, time-cycled breaths from the ventilator. The smaller “flat-top” pressure curves represent patient-triggered, pressure-limited (pressure-supported) spontaneous patient breaths. Pressure-supported spontaneous breaths between patient-triggered ventilator breaths are depicted in the figure. This is intermittent mandatory ventilation, and, more specifically, it is pressure-controlled intermittent mandatory ventilation because the ventilator breaths are pressure limited.

PTS:	1	REF:	Page 405

82.	A patient being ventilated in the volume-controlled continuous mandatory ventilation mode had a measured tidal volume of 620 mL for the breath shown in the pressure-time scalar in the figure. What is the static compliance?

| a. | 18 mL/cm H2O | c. | 31 mL/cm H2O |
| --- | --- | --- | --- |
| b. | 21 mL/cm H2O | d. | 62 mL/cm H2O |

ANS:	C

Static lung compliance = Tidal volume ÷ (Plateau pressure – Positive end-expiratory pressure) = 620 mL ÷ (30 cm H2O – 10 cm H2O) = 31 mL/cm H2O.

PTS:	1	REF:	Page 390

83.	Obstructive sleep apnea is most commonly treated with which of the following?

| a. | Pressure-support ventilation | c. | Continuous mandatory ventilation |
| --- | --- | --- | --- |
| b. | Pressure-controlled ventilation | d. | Bilevel positive airway pressure |

ANS:	D

Bilevel positive airway pressure is commonly used to treat obstructive sleep apnea.

PTS:	1	REF:	Page 390

84.	Which of the following modes is used for noninvasive positive-pressure ventilation?

| a. | Pressure-support ventilation | c. | Continuous positive airway pressure |
| --- | --- | --- | --- |
| b. | Proportional assist ventilation | d. | Bilevel positive airway pressure |

ANS:	D

Bilevel positive airway pressure is similar to continuous positive airway pressure in that both breath delivery techniques provide positive pressure during inspiration and expiration. With bilevel positive airway pressure, the inspiratory positive airway pressure is higher than the expiratory positive airway pressure. Because bilevel positive airway pressure can be either patient triggered or time triggered, it can be used in the hospital setting for noninvasive ventilation.

PTS:	1	REF:	Page 409, Box 12-17

85.	What is the total cycle time for a frequency of 14 breaths/min and a ratio of inspiratory time to expiratory time of 1:3?

| a. | 1.07 second | c. | 3.21 seconds |
| --- | --- | --- | --- |
| b. | 2.33 seconds | d. | 4.29 seconds |

ANS:	D

Total cycle time = 60 s ÷ frequency = 60 ÷ 14 breaths/min = 4.29 s.

PTS:	1	REF:	Page 395, Box 12-14

86.	What is the inspiratory time for a frequency of 30 breaths/min and an inspiratory time-to-expiratory time ratio of 1:3?

| a. | 0.13 second | c. | 0.75 second |
| --- | --- | --- | --- |
| b. | 0.50 second | d. | 2.00 seconds |

ANS:	B

Total cycle time = 60 s ÷ frequency = 60 ÷ 30 = 2 s. There are four equal parts to the total cycle time (add the “I” from the inspiratory time-to-expiratory time ratio to the “E” from the inspiratory time-to-expiratory time ratio, or 1 + 3 = 4). Inspiratory time = Total cycle time ÷ (I + E) = 2 ÷ 4 = 0.5 s.

PTS:	1	REF:	Page 395, Box 12-14

87.	What is the inspiratory time for a frequency of 20 breaths/min and an inspiratory time-to-expiratory time ratio of 1:3?

| a. | 0.25 second | c. | 1.00 second |
| --- | --- | --- | --- |
| b. | 0.75 second | d. | 1.33 seconds |

ANS:	B

Total cycle time = 60 s ÷ frequency = 60 ÷ 20 = 3 s. There are four equal parts to the total cycle time (add the “I” from the inspiratory time-to-expiratory time ratio to the “E” from the inspiratory time-to-expiratory time ratio, or 1 + 3 = 4). Inspiratory time = Total cycle time ÷ (I + E) = 3 ÷ 4 = 0.75 s.

PTS:	1	REF:	Page 395, Box 12-14

88.	What is the inspiratory time-to-expiratory time ratio given the following data: frequency of 12 breaths/min; inspiratory time of 0.90 second?

| a. | 1:0.22 | c. | 1:4.6 |
| --- | --- | --- | --- |
| b. | 1:2 | d. | 1:6.4 |

ANS:	C

Inspiratory time-to-expiratory time ratio = Inspiratory time/Expiratory time and Expiratory time = Total cycle time – Inspiratory time. Inspiratory time-to-expiratory time ratio = Inspiratory time/Total cycle time – Inspiratory time. Total cycle time = 60 ÷ 12 = 5 s. Inspiratory time-to-expiratory time ratio = 0.9/5 – 0.9 = 0.9/4.1. Reduce the inspiratory time-to-expiratory time ratio to its simplest form; divide the numerator and the denominator by inspiratory time. 0.9 ÷ 0.9/4.1 ÷ 0.9 = 1:4.56, or 1:4.6.

PTS:	1	REF:	Page 395, Box 12-14

89.	What is the inspiratory time for a tidal volume of 600 mL and a flow rate of 60 L/min?

| a. | 0.1 second | c. | 1.0 second |
| --- | --- | --- | --- |
| b. | 0.6 second | d. | 6.0 seconds |

ANS:	B

Inspiratory time = Tidal volume (in liters) ÷ Flow rate (in liters/second) = 0.6 L ÷ 1 L/s = 0.6 s.

PTS:	1	REF:	Page 395, Box 12-14

90.	What is the expiratory time when the tidal volume is 700 mL, the flow rate is 50 L/min, and the inspiratory time-to-expiratory time ratio is 1:3?

| a. | 0.42 second | c. | 2.52 seconds |
| --- | --- | --- | --- |
| b. | 0.70 second | d. | 4.67 seconds |

ANS:	C

Inspiratory time = Tidal volume (in liters) ÷ Flow rate (in liters/second) = 0.7 L ÷ 0.83 L/s = 0.84 s. Expiratory time = 3  Inspiratory time = 3  0.84 = 2.52 s.

PTS:	1	REF:	Page 395, Box 12-14

91.	What is the inspiratory time-to-expiratory time ratio for the following ventilator settings: a minute volume of 8 L/min, a respiratory rate of 10 breaths/min, and a flow rate of 40 L/min?

| a. | 1:2 | c. | 1:4 |
| --- | --- | --- | --- |
| b. | 1:3 | d. | 1:5 |

ANS:	C

40 L/min = 40 ÷ 60 L/s = 0.67 L/s. Tidal volume = VE ÷ f = 8 L/min ÷ 10 = 0.8 L. Inspiratory time = Tidal volume (in liters) ÷ Flow rate (in liters/second) = 0.8 L ÷ 0.67 L/s = 1.19 s.

Total cycle time = 60 ÷ 10 = 6 s. Expiratory time = Total cycle time – Inspiratory time = 6 s – 1.19 s = 4.81 s.

Reduce the inspiratory time-to-expiratory time ratio to its simplest form; divide the numerator and the denominator by inspiratory time.

Inspiratory time/Expiratory time = 1.19 ÷ 1.19/4.81 ÷ 1.19 = 1:4.

PTS:	1	REF:	Page 395, Box 12-14

92.	The pressure-versus-time curve in the figure represents which of the following ventilator modes?

| a. | Airway pressure-release ventilation |
| --- | --- |
| b. | Pressure-support ventilation with continuous positive airway pressure |
| c. | Pressure-controlled intermittent mandatory ventilation plus pressure support |
| d. | Volume-controlled intermittent mandatory ventilation plus pressure support |

ANS:	A

Airway pressure-release ventilation provides two levels of continuous positive airway pressure while allowing the patient to breathe spontaneously at both continuous positive airway pressure levels. The sine waveforms at both continuous positive airway pressure levels represent the spontaneous breathing of the patient. The higher continuous positive airway pressure level is usually held longer than the lower continuous positive airway pressure level—hence the term airway pressure release.

PTS:	1	REF:	Page 403, Figure 12-51

93.	In the figure, the pressure-time curve that represents pressure-controlled intermittent mandatory ventilation plus pressure support is which of the following?

| a. | A | c. | C |
| --- | --- | --- | --- |
| b. | B | d. | D |

ANS:	C

Graph C shows two pressure-limited breaths at two different pressures. The higher pressure is the pressure-limited ventilator breath, and the lower pressure is the pressure-limited spontaneous breath (or pressure-supported breath). Graph A shows volume-controlled continuous mandatory ventilation, graph B shows airway pressure-release ventilation, and graph D shows volume-controlled intermittent mandatory ventilation plus pressure support.

PTS:	1	REF:	Page 407

94.	The statement “The baseline pressure is elevated but is periodically released to a lower level for a very brief period, and the patient is allowed to breathe spontaneously at each level” describes which of the following ventilator modes?

| a. | Bilevel positive airway pressure |
| --- | --- |
| b. | Pressure-controlled inverse-ratio ventilation |
| c. | Airway pressure-release ventilation |
| d. | Mandatory minute ventilation |

ANS:	C

Airway pressure-release ventilation provides two levels of continuous positive airway pressure and allows for spontaneous breathing at both levels. The ventilator is time triggered, pressure limited, and time cycled. The low time is usually approximately 0.2 to 1 second.

PTS:	1	REF:	Page 403

95.	Which of the following ventilator modes allow patients to perform part of the work of breathing?

1. Volume-support ventilation

2. Mandatory minute ventilation

3. Continuous mandatory ventilation

4. Proportional assist ventilation

| a. | 1 and 3 | c. | 1, 2, and 3 |
| --- | --- | --- | --- |
| b. | 2 and 4 | d. | 1, 2, and 4 |

ANS:	D

Volume-support ventilation on the Maquet Servo 300 is purely assist mode and is basically a form of pressure-support ventilation except that the volume delivery is targeted over several breaths. Mandatory minute ventilation is a closed-loop form of volume-targeted or pressure-targeted ventilation used in patients who can perform part of the work of breathing and are progressing toward weaning from mechanical ventilation. Proportional assist ventilation is a method of assisting spontaneous ventilation in which the practitioner adjusts the amount of the work of breathing assumed by the ventilator.

PTS:	1	REF:	Page 398

96.	During pressure-targeted ventilation with volume guaranteed for every breath, the set volume is not met. Which of the following phase variables would cycle that breath?

| a. | Time | c. | Volume |
| --- | --- | --- | --- |
| b. | Flow | d. | Pressure |

ANS:	C

If the volume is not delivered by the time flow decreases to its set value during pressure-targeted ventilation with volume guaranteed for every breath, the flow continues at the amount set on the flow control until the set tidal volume is delivered. Inspiration becomes volume cycled.

PTS:	1	REF:	Page 389

97.	Which of the following ventilator modes is based on the equation of motion and delivers pressure, flow, and volume dependent on the patient’s inspiratory effort?

| a. | Proportional assist ventilation |
| --- | --- |
| b. | Mandatory minute ventilation |
| c. | Airway pressure-release ventilation |
| d. | Pressure-controlled continuous mandatory ventilation |

ANS:	A

The operation of proportional assist ventilation is based on the equation of motion. The amount of pressure generated by the patient’s own respiratory muscles is used as an index of inspiration effort.

PTS:	1	REF:	Page 409, Box 12-17

98.	Which of the following ventilator modes is represented by the pressure-time scalar in the figure?

| a. | Volume-controlled intermittent mandatory ventilation with positive end-expiratory pressure |
| --- | --- |
| b. | Pressure-controlled intermittent mandatory ventilation with positive end-expiratory pressure |
| c. | Pressure-augmented, volume-guaranteed breaths |
| d. | Proportional assist ventilation |

ANS:	B

The pressure-time waveform is raised from the zero line; this means that there is positive end-expiratory pressure. The flat pressure wave is indicative of pressure-limited breaths. The higher level represents the ventilator pressure-limited breath, and the lower level waveform represents the pressure-supported breaths.

PTS:	1	REF:	Page 407, Figure 11-54

99.	Which of the following modes of ventilation requires the use of a special nasogastric tube that is fitted with an electrode array?

| a. | High-frequency oscillatory ventilation | c. | High-frequency jet ventilation |
| --- | --- | --- | --- |
| b. | Neurally adjusted ventilatory assist | d. | Adaptive support ventilation |

ANS:	B

Neurally adjusted ventilatory assist is a mode of ventilation based on neural respiratory input. It relies on detection of the electrical activity of the diaphragm to control ventilator function through a special nasogastric tube that is fitted with an electrode array.

PTS:	1	REF:	Page 411

100.	The added clinical benefit of secretion removal is facilitated with which high-frequency mode of ventilation?

| a. | Jet | c. | Percussive |
| --- | --- | --- | --- |
| b. | Flow | d. | Oscillatory |

ANS:	C

The clinical benefits of high-frequency percussive ventilation may include facilitation of secretion removal.

PTS:	1	REF:	Page 414

101.	Active expiration occurs in which ventilator mode?

| a. | Proportional assist ventilation |
| --- | --- |
| b. | High-frequency jet ventilation |
| c. | Airway pressure-release ventilation |
| d. | High-frequency oscillatory ventilation |

ANS:	D

With high-frequency oscillatory ventilation, pressure is positive in the airway during the inspiratory phase and negative during the expiratory phase. Both inspiration and expiration are active.

PTS:	1	REF:	Page 379

102.	Which of the following types of high-frequency ventilation involves the delivery of pulsations under pressure through a small lumen at rates of 1.7 to 10 Hz?

| a. | Jet | c. | Positive-pressure |
| --- | --- | --- | --- |
| b. | Oscillatory | d. | Flow interruption |

ANS:	A

High-frequency jet ventilation operates by passing gas from a high-pressure source through a variable regulator that reduces the pressure to the desired working level. The gas passes through a device, usually a solenoid or a fluidic valve, which governs the amount and duration of flow. The gas jet is delivered through a specially made triple-lumen endotracheal tube, which is similar to conventional endotracheal tubes except that two additional small lines are added. One line is for delivering jet ventilation, and the other is for monitoring distal airway pressures.

PTS:	1	REF:	Page 412

103.	The movement of gases from one area of the lungs to another because of differences in the compliance and resistance of various lung regions is known as:

| a. | Pendelluft | c. | Taylor dispersion |
| --- | --- | --- | --- |
| b. | Gas streaming | d. | Molecular diffusion |

ANS:	A

Pendelluft is defined as the movement of gases from one area of the lungs to another as a result of differences in the compliance and resistance of various lung regions; this is also called out-of-phase ventilation.

PTS:	1	REF:	Page 415

104.	The erratic pattern of eddies and streams caused by the turbulent flow of high-velocity gases moving through small airways and their bifurcations is caused by which of the following?

| a. | Pendelluft | c. | Taylor dispersion |
| --- | --- | --- | --- |
| b. | Gas streaming | d. | Molecular diffusion |

ANS:	C

Taylor dispersion is the enhanced mixing of gases associated with the turbulent flow of high-velocity gases moving through small airways and their bifurcations. Taylor dispersion occurs when two gas streams meet, creating an erratic pattern of eddies and streams.

PTS:	1	REF:	Page 415

105.	The tidal volume of a breath from a high-frequency jet ventilator depends on which of the following factors?

| a. | Pulsation pressure | c. | Frequency of the driving pressure |
| --- | --- | --- | --- |
| b. | Amount of bias flow | d. | Jet orifice size |

ANS:	D

The tidal volume of a breath during high-frequency jet ventilation depends on four basic factors: (1) the length of the pulsation, (2) the amplitude or driving pressure of the jet, (3) the jet orifice size, and (4) the patient’s lung characteristics.

PTS:	1	REF:	Page 415

106.	Patient discomfort from rapid gas flow during pressure-targeted ventilation may be alleviated by which of the following adjustments to the ventilator?

| a. | Decreasing expiratory time | c. | Increasing inspiratory time |
| --- | --- | --- | --- |
| b. | Decreasing total cycle time | d. | Increasing inspiratory rise time |

ANS:	D

Adding or increasing inspiratory rise time changes the rate of gas delivery at the very beginning of inspiration. This tapers the rise of pressure and flow delivery until they reach their maximum value.

PTS:	1	REF:	Page 385

107.	The statement “The breaths are pressure-limited, but the pressure limit can be adjusted between breaths to achieve a preset tidal volume” describes which of the following modes of ventilation?

| a. | Dual mode | c. | Adaptive mode |
| --- | --- | --- | --- |
| b. | Servo mode | d. | Knowledge-based mode |

ANS:	C

This statement describes an adaptive mode because the ventilator has one set-point and automatically adjusts the set-point to achieve another set-point as the patient’s condition changes.

PTS:	1	REF:	Page 412