Chapter 15: Transport, Home Care, and Noninvasive Ventilatory Devices

MULTIPLE CHOICE

1.	The pneumatic control logic of the Airon pNeuton requires a flow of _____ L/min to operate.

| a. | 1 | c. | 3 |
| --- | --- | --- | --- |
| b. | 2 | d. | 4 |

ANS:	C

The internal control logic requires a flow of 3 L/min. The additional 1 L/min on the Airon pNeuton model A is to power the patient-disconnect alarm.

PTS:	1	REF:	Page 515

2.	What alarm(s) is (are) on the Airon pNeuton ventilator?

1. Low positive end-expiratory pressure/continuous positive airway pressure

2. Patient disconnect

3. High patient pressure

4. Low oxygen inlet pressure

| a. | 2 only | c. | 2 and 4 |
| --- | --- | --- | --- |
| b. | 1 and 4 | d. | 1, 2, 3, and 4 |

ANS:	C

The pNeuton has two functional alarms that are always active: a low oxygen inlet pressure alarm that activates at pressures less than 30 psi and a patient-disconnect alarm that sounds when a disconnect situation is sensed by the ventilator.

PTS:	1	REF:	Page 516

3.	Which of the following modes of operation are offered by the Airon pNeuton?

1. Volume-controlled intermittent mandatory ventilation

2. Pressure-controlled intermittent mandatory ventilation

3. Continuous positive airway pressure

4. Pressure-regulated volume control

| a. | 1, 2, and 3 | c. | 1, 3, and 4 |
| --- | --- | --- | --- |
| b. | 2, 3, and 4 | d. | 2 and 3 |

ANS:	A

The pNeuton offers nonsynchronized volume-controlled intermittent mandatory ventilation or pressure-controlled intermittent mandatory ventilation and positive end-expiratory pressure/continuous positive airway pressure.

PTS:	1	REF:	Pages 516-517

4.	The bias flow on the Airon pNeuton is _____ L/min.

| a. | 3 | c. | 8 |
| --- | --- | --- | --- |
| b. | 5 | d. | 10 |

ANS:	D

The pneumatic system of the pNeuton provides a continuous flow of gas through the circuit of approximately 10 L/min to meet the inspiratory demand during spontaneous breaths.

PTS:	1	REF:	Page 517

5.	The internal direct current battery of the Bio-Med Crossvent operates for approximately ____ hours.

| a. | 2 | c. | 6 |
| --- | --- | --- | --- |
| b. | 4 | d. | 8 |

ANS:	C

The internal direct current battery is operational for approximately 6 hours when fully charged.

PTS:	1	REF:	Page 517

6.	Which of the following statements is true concerning the Bio-Med Crossvent?

| a. | The battery duration is 4 hours when fully charged. |
| --- | --- |
| b. | None of the Crossvent models can ventilate infants. |
| c. | Noninvasive ventilation is one of the modes of ventilation. |
| d. | Volume- and pressure-controlled assist/control and synchronized intermittent mandatory ventilation are available. |

ANS:	D

Both volume-targeted and pressure-targeted assist/control and synchronized intermittent mandatory ventilation are available on this ventilator. The battery duration is 6 hours when fully charged. Infant ventilation is offered by several of the Crossvent models. Noninvasive ventilation is not offered by this ventilator.

PTS:	1	REF:	Page 517

7.	During transport, an infant who is using a Bio-Med MVP-10 unit in the intermittent mandatory ventilation mode with 60% fractional inspired oxygen requires 3 L/min of oxygen and 2 L/min of air to maintain the fractional inspired oxygen and desired peak pressure. An E cylinder with 2200 pounds per square inch gauge of oxygen would last _____ minutes.

| a. | 88.8 | c. | 207.3 |
| --- | --- | --- | --- |
| b. | 155.5 | d. | 314.3 |

ANS:	A

Pressure of gas supply = tank pressure  conversion factor = 2200 pounds per square inch gauge  0.28 = 622 L. 622 L ÷ (3 L/min for the oxygen + 4 L/min for the fluidic logic) = 622 ÷ 7 = 88.8 minutes.

PTS:	1	REF:	Page 521, Clinical Scenario 15-3

8.	A respiratory rate of 40 breaths/min is required to ventilate an infant with an MVP-10. The inspiratory and expiratory time settings that would provide 40 breaths/min are which of the following?

Inspiratory Time	Expiratory Time

| a. | 0.25 s 0.9 s | c. | 0.75 s 2 s |
| --- | --- | --- | --- |
| b. | 0.5 s 1 s | d. | 0.95 s 1.75 s |

ANS:	B

The total cycle time for 40 breaths/min is 60 ÷ 40 = 1.5 s. Inspiratory time plus expiratory time must equal 1.5 s. The only combination that adds up to 1.5 s is inspiratory time = 0.5 s and expiratory time = 1 s.

PTS:	1	REF:	Page 521

9.	What modes are available on the MVP-10?

1. Pressure-controlled assist/control

2. Continuous positive airway pressure

3. Pressure-regulated volume control

4. Nonsynchronous intermittent mandatory ventilation

| a. | 1 and 3 | c. | 1 and 4 |
| --- | --- | --- | --- |
| b. | 2 and 4 | d. | 2 and 3 |

ANS:	B

The MVP-10 has two modes of ventilation, nonsynchronized intermittent mandatory ventilation and continuous positive airway pressure.

PTS:	1	REF:	Pages 522-523

10.	The operation time for a Dräger Oxylog 3000 with a lithium ion battery is ___ hours.

| a. | 3 | c. | 5 |
| --- | --- | --- | --- |
| b. | 4 | d. | 6 |

ANS:	B

The Oxylog 3000 is functional for 4 hours with the lithium ion battery.

PTS:	1	REF:	Page 523, Table 15-6

11.	How is the minute ventilation high alarm (MVhigh) set on the Dräger Oxylog 3000?

| a. | Operator-set at 2 L/min above the set minute ventilation |
| --- | --- |
| b. | Operator-set at 3 L/min above the measured minute ventilation |
| c. | Automatically set at 3 L/min above the set minute ventilation |
| d. | Automatically set at 2 L/min above the measured minute ventilation |

ANS:	D

The Oxylog 3000 has multiple built-in alarms to help support quick start-up of the ventilator. The MVhigh alarm is set at the measured minute ventilation +2 L/min.

PTS:	1	REF:	Page 525

12.	The Dräger Oxylog 3000 can provide which of the following modes of operation?

1. Pressure-controlled ventilation+/pressure support

2. Continuous mandatory ventilation/continuous mandatory ventilationassist

3. Pressure-regulated volume control

4. Control

| a. | 1 and 2 | c. | 3 and 4 |
| --- | --- | --- | --- |
| b. | 2 and 3 | d. | 1 and 4 |

ANS:	A

The Oxylog 3000 can provide continuous mandatory ventilation/continuous mandatory ventilationassist, pressure-controlled ventilation+/pressure support, continuous positive airway pressure/pressure support, synchronized intermittent mandatory ventilation/pressure support, and apnea backup.

PTS:	1	REF:	Page 525

13.	During a patient transport with the Impact Uni-Vent 750, the respiratory therapist hears a whistling sound. What is the most likely cause of this whistling?

| a. | The “PEEP not set” alarm is currently active. |
| --- | --- |
| b. | The “battery low/fail” alarm is being violated. |
| c. | The patient flow is greater than the set flow. |
| d. | The demand valve has opened, and gas is passing through. |

ANS:	C

The patient valve serves the function of an exhalation valve and contains the antiasphyxia valve. If the flow setting is insufficient for the inspiratory demands of the patient, the antiasphyxia valve allows room air to be entrained for additional flow. When air entrainment occurs, the valve creates a whistling sound.

PTS:	1	REF:	Page 529

14.	The estimated internal direct current battery on the Uni-Vent 750 can provide ____ hours of operation.

| a. | 3 | c. | 9 |
| --- | --- | --- | --- |
| b. | 6 | d. | 12 |

ANS:	C

The internal battery has an estimated life of 9 hours.

PTS:	1	REF:	Page 527

15.	Which of the following statements is true concerning the antiasphyxia valve of the Uni-Vent 750?

| a. | The valve opens with a –1 cm H2O pressure. |
| --- | --- |
| b. | The valve opens in the synchronized intermittent mandatory ventilation mode for spontaneous breathing. |
| c. | The valve can open in the control mode for spontaneous breathing. |
| d. | The valve functions as the ventilator’s demand valve for breath triggering. |

ANS:	C

If the flow setting is insufficient for the inspiratory demands of the patient, the antiasphyxia valve allows room air to be entrained for additional flow. This allows the patient to breathe spontaneously between the mandatory breaths in the control mode.

PTS:	1	REF:	Page 529

16.	Pressure plateau ventilation on the Impact Uni-Vent Eagle 754 is most similar to which of the following modes of ventilation?

| a. | Airway pressure release ventilation |
| --- | --- |
| b. | Pressure-regulated volume control |
| c. | Pressure-controlled continuous mandatory ventilation |
| d. | DuoPAP |

ANS:	B

The pressure plateau ventilation mode on the Eagle 754 limits the pressure allowed in the circuit and airways but does not end inspiration until the desired tidal volume is delivered. The difference is that it is the inspiratory time that changes to ensure the volume delivery.

PTS:	1	REF:	Page 530

17.	Which of the following is true concerning the difference between the Uni-Vent 750 and the Eagle 754?

| a. | The Eagle 754 can be used for noninvasive ventilation. |
| --- | --- |
| b. | The Eagle 754 has a built-in compressor with a blender. |
| c. | The 750 is listed in the Strategic National Stockpile. |
| d. | The 750 has a gas consumption of 3 L/min for its control logic. |

ANS:	B

The Eagle 754 has an internal compressor with a blender, and it is listed in the Strategic National Stockpile for deployment. Neither the 750 nor the 754 consumes gas, and neither can be used for noninvasive ventilation.

PTS:	1	REF:	Page 531

18.	The Percussionaire Bronchotron TXP ventilator requires which of the following power sources for conventional operation?

1. 9-V battery

2. Alternating current 120 V

3. 4 AA batteries

4. 30-psi gas source

| a. | 4 only | c. | 1, 3, and 4 |
| --- | --- | --- | --- |
| b. | 2 and 4 | d. | 1, 2, and 3 |

ANS:	C

The Bronchotron is a pneumatic ventilator, requiring only a gas source with a minimum of 30 psi to operate conventionally as well as 4 AA batteries for the frequency counter and a 9-V battery for the low-pressure alarm.

PTS:	1	REF:	Page 536

19.	The amount of gas consumed by the Percussionaire Bronchotron TXP ventilator is ____ L/min.

| a. | 2 to 4 | c. | 6 to 10 |
| --- | --- | --- | --- |
| b. | 3 to 8 | d. | 8 to 12 |

ANS:	D

The Bronchotron TXP requires 8 to 12 L/min for the logic flow to operate.

PTS:	1	REF:	Page 536, Table 15-12

20.	Increasing the amplitude control on the Percussionaire Bronchotron TXP Ventilator would:

| a. | Increase the fractional inspired oxygen |
| --- | --- |
| b. | Decrease the frequency |
| c. | Increase carbon dioxide removal |
| d. | Increase air entrainment |

ANS:	C

The amplitude knob determines the amplitude of the high-frequency breaths. Increasing the amplitude would normally increase the removal of carbon dioxide.

PTS:	1	REF:	Page 538

21.	The inspiratory-to-expiratory ratio would be 1:1.5 for _____ breaths/min on the Percussionaire Bronchotron TXP Ventilator.

| a. | 155 | c. | 275 |
| --- | --- | --- | --- |
| b. | 200 | d. | 500 |

ANS:	C

A rate of 250 to 400 breaths/min would have an automatic inspiratory-to-expiratory ratio of 1:1.5.

PTS:	1	REF:	Page 538

22.	The large external battery that is used with the Pulmonetic Systems LTV 1000 ventilator lasts approximately ____ hours when fully charged.

| a. | 4 | c. | 8 |
| --- | --- | --- | --- |
| b. | 6 | d. | 10 |

ANS:	C

The fully charged large external battery can provide up to approximately 8 hours of power.

PTS:	1	REF:	Page 540

23.	The alarm sound on the Pulmonetic Systems LTV 1000 ventilator is generated out of which of the following structures in the figure?

| a. | (1) | c. | (3) |
| --- | --- | --- | --- |
| b. | (2) | d. | (4) |

ANS:	A

The right side of the Pulmonetic Systems LTV 1000 contains the following: (1) a small opening for the alarm sound that should never be covered, (2) a 22-mm connector for the main inspiratory flow line, (3) an external connection for the exhalation valve, and (4) two flow-transducer connectors that attach to the patient Y-adapter.

PTS:	1	REF:	Page 540

24.	Using the figure, what amount of oxygen is necessary to be bled into a Pulmonetic Systems LTV 1000, with a minute volume of 10 L/min, to achieve a fractional inspired oxygen level of 70%?

| a. | 5 L/min | c. | 10 L/min |
| --- | --- | --- | --- |
| b. | 8 L/min | d. | 15 L/min |

ANS:	C

Locate the minute volume of 10 L/min on the right side of the chart. Follow the 10 L/min line down until it intersects with the 0.70 vertical line. Follow the vertical line left to read the input oxygen flow (L/min).

PTS:	1	REF:	Page 543, Box 15-12

25.	During volume-targeted assist/control with a Pulmonetic Systems LTV 1000, the tidal volume is set at 800 mL, and inspiratory time is set at 1 s. Inspiration will end when which of the following is reached?

| a. | Tidal volume of 800 mL | c. | Set pressure limit |
| --- | --- | --- | --- |
| b. | Inspiratory time of 1 s | d. | Flow rate of 24 L/min |

ANS:	D

During volume-targeted breaths, the peak flow at the beginning of inspiration is calculated by the ventilator so that the tidal volume can be delivered during the set inspiratory time, and inspiration ends when flow decreases to 50% of the peak or 10 L/min, whichever is highest. The calculated flow rate is 0.8 L ÷ 1 s, or 0.8 L ÷ 1/60 of a second. Flow rate = 48 L/min; 50% of 48 L/min is 24 L/min.

PTS:	1	REF:	Page 544, Box 15-14

26.	Which of the following statements concerning the Pulmonetic Systems LTV 1000 ventilator is true?

| a. | An external positive end-expiratory pressure valve must be added. |
| --- | --- |
| b. | There is a backup pressure trigger of –3 cm H2O. |
| c. | The square-flow waveform is automatically set. |
| d. | Volume-targeted breaths are the only breath type available. |

ANS:	B

In addition to flow triggering, there is a backup pressure-sensing trigger that pressure triggers a breath when the airway pressure drops below –3 cm H2O.

PTS:	1	REF:	Page 543

27.	On the Smiths Medical Pneupac ventiPAC ventilator, which of the following would provide a 15 breaths/min respiratory rate?

| a. | Set inspiratory-to-expiratory ratio at 1:4 |
| --- | --- |
| b. | Set the rate control = 15 breaths/min |
| c. | Set inspiratory time at 1.5 s and expiratory time at 2.5 s |
| d. | Set flow rate control = 30 L/min, tidal volume = 500 mL, and expiratory time = 3 seconds |

ANS:	C

Respiratory rate is set by manipulation of the inspiratory and expiratory time control knobs. Inspiratory time is adjustable from 0.5 to 2 seconds. Expiratory time is adjustable from 0.6 to 6 seconds.

PTS:	1	REF:	Page 559

28.	What inspiratory flow and inspiratory time would produce a tidal volume of 750 mL on the Smiths Medical Pneupac ventiPAC ventilator?

Inspiratory flow	Inspiratory time

1. 0.33 L/s	1 s

2. 0.5 L/s	1.5 s

3. 0.75 L/s	1 s

4. 0.45 L/s	2 s

| a. | 3 only | c. | 1 and 4 |
| --- | --- | --- | --- |
| b. | 2 and 3 | d. | 1, 2, and 4 |

ANS:	B

Tidal volume (L) = Inspiratory flow (L/s)  Inspiratory time (s). 0.5 L/s  1.5 s = 0.75 L, or 750 mL, and 0.75 L/s  1 s = 0.75 L, or 750 mL.

PTS:	1	REF:	Page 558

29.	Which transport ventilator does not require an external blender to regulate precisely the fractional inspired oxygen?

| a. | Uni-Vent Eagle 754 | c. | Bio-Med Crossvent |
| --- | --- | --- | --- |
| b. | Bio-Med MVP-10 | d. | Airon pNeuton |

ANS:	A

The Eagle 754 has a built-in compressor with a blender.

PTS:	1	REF:	Page 530

30.	Which of the following is a magnetic resonance imaging–compatible transport ventilator?

| a. | Bio-Med MVP-10 | c. | Uni-Vent Eagle 754 |
| --- | --- | --- | --- |
| b. | Dräger Oxylog 3000 | d. | Bio-Med Crossvent |

ANS:	A

The Bio-Med MVP-10 has a magnetic resonance imaging–compatible version.

PTS:	1	REF:	Page 515

31.	Which of the following is a transport ventilator that can deliver high-frequency ventilation?

| a. | Percussionaire Bronchotron TXP | c. | Pulmonetic Systems LTV 1200 |
| --- | --- | --- | --- |
| b. | Impact Uni-Vent Eagle 754 | d. | Dräger Oxylog 3000 |

ANS:	A

The Bronchotron TXP is a high-frequency flow-interrupter ventilator capable of frequencies of 700 breaths/min.

PTS:	1	REF:	Page 536, Table 15-12

32.	The transport ventilator that can be used on newborns is the:

| a. | Airon pNeuton | c. | Bio-Med Crossvent 4 |
| --- | --- | --- | --- |
| b. | Uni-Vent Eagle 754 | d. | Pulmonetic Systems LTV 1000 |

ANS:	C

The Bio-Med Crossvent 4 and 2+ can be used for neonatal ventilation.

PTS:	1	REF:	Page 518, Table 15-2

33.	The Dräger Carinahome ventilator has which of the following ventilator modes?

1. Continuous positive airway pressure

2. Volume-controlled ventilation-pressure regulated+

3. Pressure-controlled ventilation+

4. Control

| a. | 1 and 4 | c. | 1, 2, and 3 |
| --- | --- | --- | --- |
| b. | 2 and 3 | d. | 2 and 4 |

ANS:	C

The Carinahome ventilator has continuous positive airway pressure, pressure support ventilation, pressure-controlled ventilation+ or pressure-controlled ventilation assist, and volume-controlled ventilation-pressure regulated+.

PTS:	1	REF:	Page 523

34.	Oxygen can be added to the Dräger Carinahome ventilator by which of the following methods?

| a. | Connecting to an external oxygen blender |
| --- | --- |
| b. | Connecting oxygen to the low-pressure input port |
| c. | Switching on the air-mix control, and setting the oxygen concentration |
| d. | Hooking up both air and oxygen high-pressure hoses to source gas |

ANS:	B

The Carina has an oxygen low-pressure input port, where supplemental oxygen can be bled in to a maximum of 10 L/min.

PTS:	1	REF:	Page 523

35.	The fractional inspired oxygen on a Dräger Carinahome ventilator is determined by which of the following?

| a. | Fractional inspired oxygen setting on the ventilator |
| --- | --- |
| b. | The position of the air-mix control on the ventilator |
| c. | The type of optional valve being used and the oxygen flow rate |
| d. | The oxygen flow rate being bled into the low-pressure input port |

ANS:	C

The Dräger Carinahome ventilator has an oxygen low-pressure input port, where supplemental oxygen can be bled in to a maximum of 10 L/min. The Carina has two optional valves that dictate the oxygen percentage delivered.

PTS:	1	REF:	Page 523

36.	Which of the following statements concerning the Newport HT50 ventilator is true?

| a. | It requires both an oxygen and an air source to operate properly. |
| --- | --- |
| b. | It can be used to ventilate patients who weigh more than 5 kg. |
| c. | The internal battery can sustain ventilator operation for 10 hours. |
| d. | The “battery empty” message is activated when there are less than 15 minutes left. |

ANS:	C

The internal battery has 10 hours of battery life when fully charged. This ventilator does not require external gas sources to perform all of its operations properly. It can be used on patients who weigh 10 kg or more, and the “battery empty” message activates when there are 30 minutes of battery life left.

PTS:	1	REF:	Page 566

37.	What is the tidal volume delivered when the Newport HT50 is set to deliver volume-controlled ventilation with an inspiratory time of 0.75 second and a flow of 0.83 L/s?

| a. | 375 mL | c. | 904 mL |
| --- | --- | --- | --- |
| b. | 623 mL | d. | 1107 mL |

ANS:	B

Tidal volume = Inspiratory time (s)  Flow (L/s) = 0.75 s  0.83 L/min = 623 mL.

PTS:	1	REF:	Page 566

38.	The apnea alarm on the Newport HT50 ventilator is triggered after how much time?

| a. | Operator-selected value | c. | Ventilator set at 20 seconds |
| --- | --- | --- | --- |
| b. | Ventilator set at 15 seconds | d. | Ventilator set at 30 seconds |

ANS:	D

When a patient breath goes undetected for 30 seconds, an apnea alarm occurs.

PTS:	1	REF:	Page 566

39.	During the transport of a patient being ventilated with a Newport HT50 ventilator, the low airway pressure/apnea indicator begins to blink red and the “check proximal line” alarm is triggered. What is the most appropriate action?

| a. | Suction the patient’s airway. | c. | Check the circuit’s integrity. |
| --- | --- | --- | --- |
| b. | Increase the pressure limit. | d. | Switch the mode of ventilation. |

ANS:	C

The “check proximal line” alarm occurs during inspiration when the proximal pressure measured is significantly different from the internal backup pressure sensor measurement taken inside the ventilator. In addition to the message, the low airway pressure/apnea indicator blinks red. Causes include disconnections, kinking of the circuit, and a water-filled proximal sensor line.

PTS:	1	REF:	Page 565

40.	Which of the following situations is most likely to activate the occlusion alarm on the Newport Medical Instruments HT50 ventilator?

| a. | Increased airway resistance | c. | Water in the proximal sensor line |
| --- | --- | --- | --- |
| b. | Secretions in the patient circuit | d. | Patient biting the endotracheal tube |

ANS:	D

The occlusion alarm activates when the airway pressure remains +15 cm H2O above the positive end-expiratory pressure setting for 3 seconds after exhalation has begun or at the end of exhalation, whichever comes first. This could occur when a patient bites the endotracheal tube because biting the tube can keep the pressure elevated for long periods.

PTS:	1	REF:	Page 565

41.	Which of the following is a transport ventilator that operates normally up to an altitude of 15,000 feet?

| a. | Airon pNeuton | c. | Impact Uni-Vent Eagle 754 |
| --- | --- | --- | --- |
| b. | Dräger Oxylog 2000 | d. | Pulmonetic Systems LTV 1000 |

ANS:	A

The pNeuton operates normally up to an altitude of 15,000 feet. Altitude does not affect pressure settings.

PTS:	1	REF:	Page 515

42.	What is the trigger variable for the Respironics BiPAP Focus?

| a. | Time | c. | Volume |
| --- | --- | --- | --- |
| b. | Flow | d. | Pressure |

ANS:	C

One of the two methods for triggering to inspiratory positive airway pressure breathing is volume. When 6 mL of volume from patient effort is inhaled from the circuit, the ventilator will volume trigger to inspiratory positive airway pressure breathing.

PTS:	1	REF:	Page 582

43.	The default value for “Ramp Start” on the Respironics BiPAP Focus is _____ cm H2O.

| a. | 2 | c. | 6 |
| --- | --- | --- | --- |
| b. | 4 | d. | 8 |

ANS:	B

The default level for the ramp pressure initiation level is 4 cm H2O.

PTS:	1	REF:	Page 584

44.	The Respironics BiPAP Synchrony is limited to a patient weighing ____ kg or more.

| a. | 15 | c. | 25 |
| --- | --- | --- | --- |
| b. | 20 | d. | 30 |

ANS:	D

The BiPAP Synchrony is designed to provide noninvasive ventilatory assistance to adult patients who weigh 30 kg or more.

PTS:	1	REF:	Page 589

45.	The low-priority alarms on the BiPAP Synchrony include which of the following?

| a. | Low battery | c. | Battery voltage too high |
| --- | --- | --- | --- |
| b. | Patient disconnect | d. | Momentary loss of power |

ANS:	D

“Momentary loss of power” is a low or yellow/constant alarm on the BiPAP Synchrony. “Low battery” is a medium-priority alarm (yellow flashing). “Patient disconnect” is a high-priority, red, flashing alarm. “Battery voltage too high” is a medium-priority (yellow flashing) alarm.

PTS:	1	REF:	Page 559, Table 15-21

46.	The two mechanisms that are used to identify and adjust baseline values in response to leaks in the Respironics BiPAP Vision (V) -60 are which of the following?

1. Tidal volume

2. Total flow rate

3. Inspiratory time

4. Expiratory flow rate

| a. | 1 and 2 | c. | 1 and 4 |
| --- | --- | --- | --- |
| b. | 3 and 4 | d. | 2 and 3 |

ANS:	C

The BiPAP Vision has a system called Auto-Trak sensitivity that uses a combination of two mechanisms to identify and modify baseline values in response to leaks. These are expiratory flow rate and tidal volume adjustments.

PTS:	1	REF:	Page 582

47.	Which of the following statements concerning the Puritan Bennett GoodKnight 425 is true?

| a. | This unit has a leak alarm and a power disconnect alarm. |
| --- | --- |
| b. | This unit can be used for noninvasive and invasive ventilation. |
| c. | If apneic events are a significant issue, the patient should not use this unit. |
| d. | Oxygen is supplied by a high-pressure hose attached to the back of the unit. |

ANS:	C

The bilevel mode supplied by this ventilator is not a timed mode. It is a spontaneous mode only. Because there are no alarms, use of this ventilator on a patient who has apnea episodes is contraindicated.

PTS:	1	REF:	Page 591

48.	The Crossvent requires a compressed gas source of _________ to ensure adequate gas flow to the patient.

| a. | 25 to 40 psi | c. | 31 to 75 psi |
| --- | --- | --- | --- |
| b. | 30 to 50 psi | d. | 40 to 90 psi |

ANS:	C

The Crossvent requires a compressed gas source of 31 to 75 psi to ensure adequate gas flow to the patient. If the gas pressure drops below 31 psi, a low pressure alarm activates.

PTS:	1	REF:	Page 517

49.	What is the maximum length of an inspiratory hold on the Oxylog 3000 plus?

| a. | 10 seconds | c. | 15 seconds |
| --- | --- | --- | --- |
| b. | 12 seconds | d. | 18 seconds |

ANS:	C

The Oxylog 3000 plus either extends the momentary breath or starts a new inspiratory breath and holds it for a maximum of 15 seconds.

PTS:	1	REF:	Page 524

50.	How long does the Oxylog 3000 plus provide 100% oxygen for suctioning?

| a. | 60 seconds | c. | 180 seconds |
| --- | --- | --- | --- |
| b. | 120 seconds | d. | 240 seconds |

ANS:	C

Depressing the soft key increases the fractional inspired oxygen to 1.0 for 3 minutes regardless of the preset fractional inspired oxygen.

PTS:	1	REF:	Page 525

51.	When using a Newport HT-50 ventilator, what is the maximum of the sum of the pressure support and positive end-expiratory pressure level?

| a. | 50 cm H2O | c. | 60 cm H2O |
| --- | --- | --- | --- |
| b. | 55 cm H2O | d. | 65 cm H2O |

ANS:	C

The ventilator does not allow a combined positive end-expiratory pressure/continuous positive airway pressure value and pressure support level of greater than 60 cm H2O.

PTS:	1	REF:	Page 564

52.	Which of the following alarms are included on a Respironics V-60?

1. High peak airway pressure alarm

2. Low inspiratory pressure alarm

3. Low end-expiratory positive airway pressure alarm

4. High tidal volume alarm

5. Low minute ventilation alarm

| a. | 1 and 3 | c. | 1, 2, 4, and 5 |
| --- | --- | --- | --- |
| b. | 1, 2, and 5 | d. | 1, 2, 3, 4, and 5 |

ANS:	C

The Respironics V-60 includes a high peak airway pressure alarm, a low inspiratory pressure alarm, a high and low tidal volume alarm, a high rate alarm, a low minute ventilation alarm, and a low inspiratory pressure delay time.

PTS:	1	REF:	Page 584

53.	Which of the following modes of ventilation on the Respironics V-60 delivers a pressure that relates to the patient’s inspiratory effort?

| a. | Proportional pressure ventilation |
| --- | --- |
| b. | Average volume-assured pressure support |
| c. | Pressure-controlled ventilation |
| d. | Continuous positive airway pressure |

ANS:	A

Proportional pressure ventilation increases the pressure as the patient’s inspiratory effort increases.

PTS:	1	REF:	Page 584

54.	Which of the following interfaces is (are) used in the application of noninvasive positive pressure ventilation?

1. Nasal mask

2. Nasal pillow

3. Oronasal mask

4. Total face mask

| a. | 1 only | c. | 1, 2, and 4 |
| --- | --- | --- | --- |
| b. | 1 and 3 | d. | 1, 2, 3, and 4 |

ANS:	D

The interfaces that are presently being used include the nasal mask, nasal pillow, oronasal mask, total face mask, and mouthpiece.

PTS:	1	REF:	Page 578